## EMDN
European Medical Device Nomenclature (EMDN)
A - Devices for administration, withdrawal and collection | B - Haematology and haemotransfusion devices | C - Cardiocirculatory system devices | D - Disinfectants, antiseptics, sterilising, agents and detergents for medical devices | F - Dialysis devices | G - Gastrointestinal devices | H - Suture devices | J - Active - implantable devices | K - Endotherapy and electrosurgical devices | L - Reusable surgical instruments | M - Devices for general and specialist dressings | N - Nervous and medullary systems devices | P - Implantable prosthetic and osteosynthesis devices | Q - Dental, ophthalmologic and ent devices | R - Respiratory and anaesthesia devices | S - Sterilisation devices (excluding CAT.D - Z) | T - Patient protective equipment and incontinence aids (excluding personal protective equipment - PPE) | U - Devices for urogenital system | V - Various medical devices | W - In vitro diagnostic medical devices | Y - Devices for persons with disabilities not included in other categories | Z - Medical equipment and related accessories, software and consumables
A01 - Needles | B01 - Blood bags and kits | C01 - Arterio - venous system devices | D01 - Aldehydes for the disinfection of medical devices | F01 - Dialysis filters | G01 - Oro - oesophageal tubes | H01 - Surgical sutures | J01 - Cardiac functionality implantable devices | K01 - Endotherapy devices | L01 - Sharp instruments, reusable | M01 - Cotton and synthetic wadding | N01 - Encephalic and peripheral nervous system devices | P01 - Facial and odontological prostheses | Q01 - Dental devices | R01 - Intubation devices | S01 - Devices for sterilisation and packaging (excluding CAT.D - Z) | T01 - Gloves (excluding personal protective equipment - PPE) | U01 - Urethral, prostatic and bladder catheters | V01 - Cutting devices, single - use | W01 - Reagents | Y05 - Abilities learning and training devices | Z11 - Bioimaging and radiotherapy instruments
A02 - Syringes | B02 - Blood filters | C02 - Arrhythmology devices | D02 - Biguanides for the disinfection of medical devices | F02 - Dialysis lines | G02 - Gastrointestinal tubes and sets | H02 - Mechanical surgical staplers | J02 - Implantable neurostimulators | K02 - Electrosurgery devices, single - use | L02 - Suture instruments, reusable | M02 - Gauzes | N02 - Spinal medullar system devices | P02 - Ent prostheses | Q02 - Ophthalmic devices | R02 - Breathing circuits and catheter mounts | S90 - Devices for sterilisation (excluding cat. D - z) - various | T02 - Protective clothing and drapes (excluding personal protective equipment - PPE) | U02 - Ureteral catheters and stents | V02 - Neonatology and paediatric devices | W02 - IVD instruments | Y06 - External prostheses and orthoses | Z12 - Instruments for functional explorations and therapeutic interventions
A03 - Tubular devices | B03 - Apheresis devices | C03 - Cardiac surgery and heart transplant devices | D03 - Chlorine derivatives for the disinfection of medical devices | F03 - Dialysis kits | G03 - Gastrointestinal endoscopy devices | H03 - Haemostasis clips | J03 - Auditory active - implantable devices | K03 - Arthroscopy devices, single - use | L03 - General surgery instruments, reusable | M03 - Bandages | N99 - Nervous and medullary system devices - other | P03 - Ocular prostheses | Q03 - Ent devices | R03 - Respiratory masks and balloons, single - use and reusable |  | T03 - Protections (excluding personal protective equipment PPE) | U03 - Devices for urethral, ureteral and nephrostomic dilation | V03 - Measurement devices | W05 - IVD generic use consumables | Y09 - Personal care aids | Z13 - Non - specific consumables for diagnostic instruments
A04 - Solution filters | B04 - Autotransfusion devices | C04 - Cardiovascular guidewires | D05 - Oxygen producers for disinfection of medical devices | F04 - Dialysis concentrates | G04 - Orally administered gastrointestinal devices | H90 - Suture devices - various | J04 - Implantable pumps |  | L04 - Abdominal surgery instruments, reusable | M04 - Special dressings |  | P05 - Oesophageal and gastrointestinal prostheses |  | R04 - Respiratory filters |  | T04 - Incontinence devices | U04 - Devices for percutaneous urinary drainage and nephrostomy catheters | V04 - Clinical use containers (non - IVD) |  | Y12 - Personal mobility devices
A05 - Mechanical infusion systems, single - use | B05 - Topical use blood components preparation devices | C05 - Cardiovascular introducer sheaths | D06 - Phenols for disinfection of medical devices | F05 - Devices for microdialysis of specific organs | G05 - Topical anorectal administered gastro - intestinal system devices |  | J05 - Brachytherapy implantable devices |  | L05 - Obstetrics and gynecology instruments, reusable | M05 - Plasters |  | P06 - Breast implants |  | R05 - Suction and dilatation systems for the respiratory system |  | T99 - Protective devices and aids for incontinence (excluding personal protective equipment - ppe) - other | U05 - Urodynamics devices | V05 - Clinical procedures kits not included in other classes |  | Y15 - Domestic activities aids
A06 - Drainage and fluids collection devices | B06 - Cellular or biological manipulation devices | C06 - Cardiovascular surgery instruments, single - use | D07 - Alcohols for the disinfection of medical devices | F90 - Dialysis devices - various | G99 - Gastrointestinal devices - other |  | J06 - Active implantable glucose monitoring systems |  | L06 - Urology instruments, reusable | M90 - Devices for general and specialist dressings - various |  | P07 - Vascular and cardiac prostheses |  | R06 - Nebulisation and humidification systems |  |  | U06 - Urological guidewires | V06 - Clinical procedure simulation devices |  | Y18 - Aids, furniture and home aids for persons with disabilities
A07 - Adapters, connectors , ramps, stopcocks, caps | B07 - Blood and blood components preservation solutions | C90 - Cardiocirculatory devices - various | D08 - Detergents for medical devices |  |  |  | J07 - Ophthalmological use active implantable devices |  | L07 - Cardiovascular surgery instruments, reusable |  |  | P08 - Urogenital prostheses |  | R07 - Bronchopulmonary endoscopy device, signal - use |  |  | U07 - Devices for the treatment of incontinence | V07 - Medical device cleaning devices, not otherwise classified |  | Y21 - Communication and information management aids
A08 - Nutrition and infusion bags and containers, signal - use | B99 - Haematology and haemotransfusion devices - other | C99 - Cardiocirculatory devices - other | D09 - Ammonium salts and associates for the disinfection of medical devices |  |  |  | J99 - Active-implantable devices - other |  | L08 - Thoracic surgery instruments, reusable |  |  | P09 - Orthopaedic prostheses, osteosynthesis device, devices for tendon and ligament synthesis |  | R90 - Respiratory and anaesthesia devices - various |  |  | U08 - Gynagecological devices | V08 - Healthcare activities supportive equipment and aids |  | Y24 - Object manipulation devices
A09 - Organ containers |  |  | D99 - Disinfectants, antiseptics, sterilising agents and detergents for medical devices - other | [5 empty] | L09 - Orthopaedic and traumatological surgery instruments, reusable |  |  | P10 - Extra - vascular support prostheses |  |  |  |  | U09 - Urogenital endoscopic devices | V09 - Fluids/Gases for clinical/therapeutic use |  | Y99 - Devices for persons with disabilities not included in other categories - other
A10 - Abdominal ostomy devices | [8 empty] | L10 - Microsurgery instruments, reusable |  |  | P90 - Implantable prosthetic and osteosynthesis devices - various |  |  |  |  | U10 - Contraceptive devices | V80 - Clinical use accessories not included in other classes
A11 - Sample collection swabs | [8 empty] | L11 - Neurosurgery and spinal surgery instruments, reusable | [7 empty] | U12 - Single - use instrumentation for urogenital system (non - endoscopic) | V90 - Various devices not included in other classes
A12 - Sample collection spatulas | [8 empty] | L12 - Laparoscopic and thoracoscopic surgery instruments, reusable | [7 empty] | U13 - Solutions and kits for bladder washing and instillation | V92 - Medical device software - not included in other classes
A99 - Devices for administration, withdrawal and collection - other | [8 empty] | L13 - robotic surgery instruments, reusable | [7 empty] | U99 - Devices for urogenital system - other
L14 - Ent instruments, reusable
L15 - Odontostomatology instruments, reusable
L16 - Diagnostic instruments, reusable
L17 - Ophthalmology instruments, reusable
L18 - Electroscopic surgery instruments, reusable
L19 - Arthroscopic surgery instruments, reusable
L20 - Thyroid surgery instruments, reusable
L21 - Clip removal surgical forceps, reusable
L22 - implantable tissues handing grasping forceps, reusable
L23 - Surgical tunnelization forceps, reusable
L24 - Dermatological surgery instruments, reusable
L25 - Plastic surgery instruments, reusable
L26 - Surgical screwdrivers, reusable
L80 - Surgical instruments, reusable - various accessories
L99 - Surgical instruments, reusable - other
Available European Medical Device Nomenclature (EMDN)
A - Devices for administration, withdrawal and collection | B - Haematology and haemotransfusion devices | J - Active - implantable devices | P - Implantable prosthetic and osteosynthesis devices | W - In vitro diagnostic medical devices | Z - Medical equipment and related accessories, software and consumables
A01 - Needles | B01 - Blood bags and kits | J02 - Implantable neurostimulators | P09 - Orthopaedic prostheses, osteosynthesis device, devices for tendon and ligament synthesis | W02 - IVD instruments | Z11 - Bioimaging and radiotherapy instruments
A02 - Syringes | B03 - Apheresis devices

## A01
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s) | Risk classification | Apply standards
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 IEC 60601-1 IEC 60601-1-2 IEC 62366-1 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | A010101 - Hypodermic needles | EU MDR, Class IIa | ISO 10993-1 ISO 7864 ISO 9626 ISO 80369-7
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 14971 | Medical devices — Application of risk management to medical devices | A010102 - Butterfly needles | EU MDR, Class IIa
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | A010103 - Needles and kits for implantable systems (port) | EU MDR, Class III
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10993-4 | Biological evaluation of medical devices Part 4: Selection of tests for interactions with blood | A010104 - Needles for vial collection | EU MDR, Class I
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | A010105 - needles for collection under vacuum | EU MDR, Class IIa
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization | A010106 - Needles and kits for reconstructive filling | N/A
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10993-11 | Biological evaluation of medical devices Part 11: Tests for systemic toxicity | A010201 - Needles and kits - histological and cytological biopsy of soft tissues | EU MDR, Class IIa
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10993-12 | Biological evaluation of medical devices Part 12: Sample preparation and reference materials | A010202 - Bone marrow biopsy needles and kits | EU MDR, Class IIa
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10993-23 | Biological evaluation of medical devices Part 23: Tests for irritation | A010203 - Cutaneous biopsy needles and kits | EU MDR, Class IIa
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10555-5 | Intravascular catheters — Sterile and single-use catheters Part 5: Over-needle peripheral catheters | A010204 - Amniocentesis and villocentesis needles and kits | EU MDR, Class IIa
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO 7864 | Sterile hypodermic needles for single use — Requirements and test methods | A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) | EU MDR, Class Ilb
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 9626 | Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods | A010206 - Cerebral biopsy needles and kits | N/A
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 23908 | Sharps injury protection — Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling | A010207 - Pulmonary biopsy needles and kits | N/A
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 80369-1 | Small-bore connectors for liquids and gases in healthcare applications Part 1: General requirements | A010301 - Spinal and epidural anaesthesia needles and kits | EU MDR, Class III
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 80369-7 | Small-bore connectors for liquids and gases in healthcare applications Part 7: Connectors for intravascular or hypodermic applications | A010302 - Plexus block needles and kits | N/A
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | IEC 60601-1 | Medical electrical equipment - ALL PARTS | A010401 - Arteriovenous fistula needles | EU MDR, Class IIa
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | IEC 60601-1-2 | Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests | A010501 - Anterior eye chamber needles | N/A
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices | A010502 - Eye irrigation needles | N/A
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 IEC 60601-1 IEC 60601-1-2 IEC 62366-1 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 8536-4 | Infusion equipment for medical use Part 4: Infusion sets for single use, gravity feed | A010503 - Peribulbar needles | N/A
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 IEC 60601-1 IEC 60601-1-2 ISO 8536-4 ISO 11607-1 ISO 11607-2 ISO 20417 IEC 63000 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 10993-7 | Biological evaluation of medical devices Part 7: Ethylene oxide sterilization residuals | A010504 - Retrobulbar needles | N/A
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices | A010505 - Endocular injection needles | N/A
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems | A010506 - Endocular aspiration cannula needles | EU MDR, Class IIa
|  |  |  | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes | A010601 - Carpule needles | EU MDR, Class IIa
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s) | ISO 20417 | Medical devices — Information to be supplied by the manufacturer | A010602 - Dental irrigation needles | EU MDR, Class I
10. Chemical, physical and biological properties |  |  |  | IEC 63000 | Amendment 1 - Technical documentation for the assessment of electrical and electronic products with respect to the restriction of hazardous substances
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to:
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 IEC 60601-1 IEC 60601-1-2 IEC 62366-1 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; | N | N/A | N/A
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(g) surface properties; and | Y | ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(h) the confirmation that the device meets any defined chemical and/or physical specifications. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
10.4. Substances
10.4.1. Design and manufacture of devices
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that:
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | Y | ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | Y | ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or | N | N/A | N/A
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. | N | N/A | N/A
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; | N | N/A | N/A
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; | N | N/A | N/A
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and | N | N/A | N/A
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. | N | N/A | N/A
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. | N | N/A | N/A
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. | N | N/A | N/A
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 20417 | A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | N | N/A | N/A
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-23 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 10993-7 ISO 11135 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. | N | N/A | N/A
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010601 - Carpule needles
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010602 - Dental irrigation needles
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 20417 | A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; | N | N/A | N/A
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; | N | N/A | N/A
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. | N | N/A | N/A
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; | N | N/A | N/A
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; | N | N/A | N/A
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. | N | N/A | N/A
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. | N | N/A | N/A
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 IEC 60601-1 IEC 60601-1-2 IEC 62366-1 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | Y | ISO 14971 IEC 60601-1 IEC 62366-1 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 IEC 60601-1 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 IEC 62366-1 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. | Y | ISO 13485 ISO 10555-5 ISO 7864 ISO 9626 ISO 23908 ISO 80369-1 ISO 80369-7 | A010101 - Hypodermic needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010106 - Needles and kits for reconstructive filling A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010106 - Needles and kits for reconstructive filling A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | Y | ISO 13485 ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 63000 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 63000 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 63000 | A010103 - 植入式系統針頭和套件（端口） A010106 - 用於重建填充的針頭和套件 A010201 - 針頭和套件 - 軟組織的組織學和細胞學活檢 A010202 - 骨髓活檢針和套件 A010203 - 皮膚切片針和套件 A010204 - 羊膜穿刺術和絨毛穿刺針和套件 A010205 - 內視鏡切片針和套件（非胃腸道設備） A010206 - 腦部活檢針和套件 A010207 - 肺活檢針和套件 A010301 - 脊髓和硬膜外麻醉針和套件 A010302 - 神經叢阻斷針和套件 A010506 - 眼內抽吸插管針 A010602 - 牙科沖洗針
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | Y | ISO 13485 ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 63000 | A010103 - 植入式系統針頭和套件（端口） A010106 - 用於重建填充的針頭和套件 A010201 - 針頭和套件 - 軟組織的組織學和細胞學活檢 A010202 - 骨髓活檢針和套件 A010203 - 皮膚切片針和套件 A010204 - 羊膜穿刺術和絨毛穿刺針和套件 A010205 - 內視鏡切片針和套件（非胃腸道設備） A010206 - 腦部活檢針和套件 A010207 - 肺活檢針和套件 A010301 - 脊髓和硬膜外麻醉針和套件 A010302 - 神經叢阻斷針和套件 A010506 - 眼內抽吸插管針 A010602 - 牙科沖洗針
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | N | N/A | N/A
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | N | N/A | N/A
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | N | N/A | N/A
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | N | N/A | N/A
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and | N | N/A | N/A
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | N | N/A | N/A
— ageing of the materials used, | N | N/A | N/A
— excess heat generated by the device, | N | N/A | N/A
— decreased accuracy of any measuring or control mechanism. | N | N/A | N/A
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | N | N/A | N/A
— the reliability of the source of energy. | N | N/A | N/A
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | N | N/A | N/A
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | N | N/A | N/A
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 62366-1 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 62366-1 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | Y | ISO 10555-5 ISO 7864 ISO 80369-1 ISO 80369-7 ISO 8536-4 | A010101 - Hypodermic needles A010102 - Butterfly needles A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010401 - Arteriovenous fistula needles A010505 - Endocular injection needles A010601 - Carpule needles
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | Y | ISO 14971 | A010101 - Hypodermic needles A010102 - Butterfly needles A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010401 - Arteriovenous fistula needles A010505 - Endocular injection needles A010601 - Carpule needles
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | Y | IEC 62366-1 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. |  | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— tissues or cells, or their derivatives, of human origin, or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | Y | IEC 62366-1 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) a declaration that the device is in a sterile condition, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) the method of sterilisation, | Y | ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) the name and address of the manufacturer, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(e) a description of the device, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(g) if the device is custom-made, the words ‘custom-made device’, | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(h) the month and year of manufacture, | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(e) the performance characteristics of the device; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | Y | ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | Y | ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— identification of any consumable components and how to replace them, | Y | ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | Y | ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | Y | ISO 14971 ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | N | N/A | N/A
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | Y | ISO 14971 ISO 11607-1 ISO 11607-2 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or | Y | IEC 62366-1 ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— information on any known restrictions to combinations of devices and equipment; | Y | IEC 62366-1 ISO 20417 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | Y | IEC 60601-1 IEC 60601-1-2 ISO 20417 IEC 63000 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | Y | IEC 60601-1 IEC 60601-1-2 ISO 20417 IEC 63000 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 ISO 20417 IEC 63000 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | ISO 14971 ISO 20417 IEC 63000 | A010103 - Needles and kits for implantable systems (port) A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010506 - Endocular aspiration cannula needles A010602 - Dental irrigation needles
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | Y | ISO 14971 ISO 10993-1 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | Y | ISO 14971 ISO 10993-1 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; | N | N/A | N/A
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | Y | ISO 14971 ISO 10993-1 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | N | N/A | N/A
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
— physical hazards such as from sharps. | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 20417 | A010101 - Hypodermic needles A010102 - Butterfly needles A010103 - Needles and kits for implantable systems (port) A010104 - Needles for vial collection A010105 - needles for collection under vacuum A010106 - Needles and kits for reconstructive filling A010201 - Needles and kits - histological and cytological biopsy of soft tissues A010202 - Bone marrow biopsy needles and kits A010203 - Cutaneous biopsy needles and kits A010204 - Amniocentesis and villocentesis needles and kits A010205 - Perendoscopic biopsy needles and kits (non-gastrointestinal apparatus) A010206 - Cerebral biopsy needles and kits A010207 - Pulmonary biopsy needles and kits A010301 - Spinal and epidural anaesthesia needles and kits A010302 - Plexus block needles and kits A010401 - Arteriovenous fistula needles A010501 - Anterior eye chamber needles A010502 - Eye irrigation needles A010503 - Peribulbar needles A010504 - Retrobulbar needles A010505 - Endocular injection needles A010506 - Endocular aspiration cannula needles A010601 - Carpule needles A010602 - Dental irrigation needles
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | N | N/A | N/A
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).

## A02
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Appl Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s) | Risk classification
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | A020101 - Loss-of-resistance syringes | EU MDR, Class I
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 14971 | Medical devices — Application of risk management to medical devices | A020102 - Infusion and irrigation syringes, single-use | EU MDR, Class I / Class IIa
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | A020104 - Syringes for injectors, single-use | EU MDR, Class IIa
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10993-4 | Biological evaluation of medical devices Part 4: Selection of tests for interactions with blood | A020105 - Blood gas analysis, syringes with needles and kits | N/A
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | A020106 - Insulin syringes, single-use | EU MDR, Class IIa
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization | A020107 - Prefilled syringes | EU MDR, Class IIa / Class III
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 10993-11 | Biological evaluation of medical devices Part 11: Tests for systemic toxicity | A020108 - Enteral feeding syringes | EU MDR, Class I / Class IIa
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 10993-23 | Biological evaluation of medical devices Part 23: Tests for irritation | A020109 - Tuberculin syringes, single-use | EU MDR, Class IIa
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 23908 | Sharps injury protection — Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling | A020201 - Reusable infusion syringes | N/A
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 7864 | Sterile hypodermic needles for single use — Requirements and test methods | A020202 - Reusable irrigation syringes | EU MDR, Class I
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO 7886-1 | Sterile hypodermic syringes for single use Part 1: Syringes for manual use | A020203 - Cartridge syringes | EU MDR, Class I / Class Ila / Class Ilb
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 7886-2 | Sterile hypodermic syringes for single use Part 2: Syringes for use with power-driven syringe pumps
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 7886-4 | Sterile hypodermic syringes for single use Part 4: Syringes with re-use prevention feature
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 80369-7 | Small-bore connectors for liquids and gases in healthcare applications Part 7: Connectors for intravascular or hypodermic applications
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 8537 | Sterile single-use syringes, with or without needle, for insulin
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | ISO 9626 | Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 10993-7 | Biological evaluation of medical devices Part 7: Ethylene oxide sterilization residuals
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | ISO 20417 | Medical devices — Information to be supplied by the manufacturer
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s)
10. Chemical, physical and biological properties
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to:
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | N | N/A | N/A
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; | N | N/A | N/A
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(g) surface properties; and | Y | ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(h) the confirmation that the device meets any defined chemical and/or physical specifications. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
10.4. Substances
10.4.1. Design and manufacture of devices
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that:
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | Y | ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | Y | ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 20417 | A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | Y | ISO 13485 ISO 14971 ISO 10993-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 10993-7 ISO 11135 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020108 - Enteral feeding syringes A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body. |  |  |  | The syringe may contain any liquid, blood or someone else, so I'm unsure if it meets the 12. and the 13. requirement?
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | Y | ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 | A020107 - Prefilled syringes | A020106
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | Y | ISO 14971 ISO 10993-1 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-23 | A020107 - Prefilled syringes
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | N | N/A | N/A
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 80369-7 ISO 8537 ISO 9626 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | N | N/A | N/A
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | Y | ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. | Y | ISO 14971 ISO 10993-1 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | N | N/A | N/A
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | N | N/A | N/A
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | N | N/A | N/A
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | N | N/A | N/A
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | N | N/A | N/A
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | N | N/A | N/A
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | N | N/A | N/A
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | N | N/A | N/A
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and | N | N/A | N/A
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | N | N/A | N/A
— ageing of the materials used, | N | N/A | N/A
— excess heat generated by the device, | N | N/A | N/A
— decreased accuracy of any measuring or control mechanism. | N | N/A | N/A
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | N | N/A | N/A
— the reliability of the source of energy. | N | N/A | N/A
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | N | N/A | N/A
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | N | N/A | N/A
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. | Y | ISO 13485 ISO 14971 IEC 62366-1 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | N | N/A | N/A
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | N | N/A | N/A
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | Y | ISO 23908 ISO 7864 ISO 7886-1 ISO 7886-2 ISO 7886-4 ISO 8537 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | Y | IEC 62366-1 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
— tissues or cells, or their derivatives, of human origin, or | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020106 - Insulin syringes, single-use A020109 - Tuberculin syringes, single-use
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | Y | IEC 62366-1 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(b) a declaration that the device is in a sterile condition, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(c) the method of sterilisation, | Y | ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(d) the name and address of the manufacturer, | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(e) a description of the device, | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(g) if the device is custom-made, the words ‘custom-made device’, | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(h) the month and year of manufacture, | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(e) the performance characteristics of the device; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | N | N/A | N/A
— identification of any consumable components and how to replace them, | N | N/A | N/A
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | N | N/A | N/A
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | N | N/A | N/A
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020105 - Blood gas analysis, syringes with needles and kits A020108 - Enteral feeding syringes A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020105 - Blood gas analysis, syringes with needles and kits A020108 - Enteral feeding syringes A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | N | N/A | N/A
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | Y | ISO 14971 ISO 11607-1 ISO 11607-2 ISO 20417 | A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020106 - Insulin syringes, single-use A020109 - Tuberculin syringes, single-use
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or | Y | IEC 62366-1 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
— information on any known restrictions to combinations of devices and equipment; | Y | IEC 62366-1 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | N | N/A | N/A
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | N | N/A | N/A
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020105 - Blood gas analysis, syringes with needles and kits A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020203 - Cartridge syringes
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | Y | ISO 14971 ISO 10993-1 ISO 20417 | A020107 - Prefilled syringes
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | Y | ISO 14971 ISO 10993-1 ISO 20417 | A020107 - Prefilled syringes
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; | Y | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | Y | ISO 14971 ISO 10993-1 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | N | N/A | N/A
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
— physical hazards such as from sharps. | Y | ISO 14971 ISO 20417 | A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020203 - Cartridge syringes
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 20417 | A020101 - Loss-of-resistance syringes A020102 - Infusion and irrigation syringes, single-use A020104 - Syringes for injectors, single-use A020105 - Blood gas analysis, syringes with needles and kits A020106 - Insulin syringes, single-use A020107 - Prefilled syringes A020108 - Enteral feeding syringes A020109 - Tuberculin syringes, single-use A020201 - Reusable infusion syringes A020202 - Reusable irrigation syringes A020203 - Cartridge syringes
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | N | N/A | N/A
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A

## B01
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s) | Risk classification
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | B010101 - Autologous blood collection bags | N/A
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 14971 | Medical devices — Application of risk management to medical devices | B010102 - Homologous donor blood collection bags and kits | EU MDR, Class Ilb
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | B010201 - Whole blood, red blood cells or plasma transfer bags and kits | EU MDR, Class Ila
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 1135-4 | Transfusion equipment for medical use Part 4: Transfusion sets for single use, gravity feed | B010202 - Platelets concentrate transfer bags and kits | EU MDR, Class Ilb / Class III
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 1135-5 | Transfusion equipment for medical use Part 5: Transfusion sets for single use with pressure infusion apparatus | B0103 - Medullary blood collection, purification, cryopreservation bags and kits | EU MDR, Class Ila
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 11737-1 | Sterilization of health care products — Microbiological methods Part 1: Determination of a population of microorganisms on products | B010401 - Placental blood collection bags and kits | N/A
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO/TS 23128:2019 | Medical devices — Transfusion set and blood bag compatibility test method | B010402 - Placental blood purification bags and kits | N/A
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 3826-1 | Plastics collapsible containers for human blood and blood components Part 1: Conventional containers | B010403 - Placental blood croypreservation bags and kits | N/A
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 3826-3 | Plastics collapsible containers for human blood and blood components Part 3: Blood bag systems with integrated features | B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | EU MDR, Class I / Class Ila
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 6710 | Single-use containers for human venous blood specimen collection
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO 80369-7 | Small-bore connectors for liquids and gases in healthcare applications Part 7: Connectors for intravascular or hypodermic applications
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 8536-4 | Infusion equipment for medical use Part 4: Infusion sets for single use, gravity feed
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 10993-7 | Biological evaluation of medical devices Part 7: Ethylene oxide sterilization residuals
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 3826-2 | Plastics collapsible containers for human blood and blood components Part 2: Graphical symbols for use on labels and instruction leaflets
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 ISO 1135-4 ISO 1135-5 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits | ISO 20417 | Medical devices — Information to be supplied by the manufacturer
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 1135-4 ISO 1135-5 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s)
10. Chemical, physical and biological properties
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to:
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10993-1 ISO 1135-4 ISO 1135-5 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | N | N/A | N/A
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 1135-4 ISO 1135-5 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; | N | N/A | N/A
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(g) surface properties; and | Y | ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(h) the confirmation that the device meets any defined chemical and/or physical specifications. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
10.4. Substances
10.4.1. Design and manufacture of devices
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that:
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | Y | ISO 14971 ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | Y | ISO 14971 ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 ISO 1135-4 ISO 1135-5 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | Y | ISO 13485 ISO 14971 ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 11737-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 11737-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 10993-7 ISO 11135 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. | N | N/A | N/A
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | Y | ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | Y | ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | N | N/A | N/A
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 1135-4 ISO 1135-5 ISO 11737-1 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | N | N/A | N/A
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | N | N/A | N/A
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | N | N/A | N/A
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | N | N/A | N/A
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | N | N/A | N/A
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | N | N/A | N/A
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | N | N/A | N/A
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | N | N/A | N/A
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | N | N/A | N/A
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | N | N/A | N/A
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and | N | N/A | N/A
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | N | N/A | N/A
— ageing of the materials used, | N | N/A | N/A
— excess heat generated by the device, | N | N/A | N/A
— decreased accuracy of any measuring or control mechanism. | N | N/A | N/A
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | N | N/A | N/A
— the reliability of the source of energy. | N | N/A | N/A
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | N | N/A | N/A
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | N | N/A | N/A
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 62366-1 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | N | N/A | N/A
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | N | N/A | N/A
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | N | N/A | N/A
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. | N | N/A | N/A
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | Y | ISO 14971 ISO 1135-4 ISO 1135-5 ISO/TS 23128:2019 ISO 3826-1 ISO 3826-3 ISO 6710 ISO 80369-7 ISO 8536-4 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | Y | ISO 14971 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | Y | IEC 62366-1 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— tissues or cells, or their derivatives, of human origin, or | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | N | N/A | N/A
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | Y | IEC 62366-1 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) a declaration that the device is in a sterile condition, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) the method of sterilisation, | Y | ISO 10993-7 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(d) the name and address of the manufacturer, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(e) a description of the device, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(g) if the device is custom-made, the words ‘custom-made device’, | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(h) the month and year of manufacture, | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | ISO 11607-1 ISO 11607-2 ISO 3826-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(e) the performance characteristics of the device; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | N | N/A | N/A
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | N | N/A | N/A
— identification of any consumable components and how to replace them, | N | N/A | N/A
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | N | N/A | N/A
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | N | N/A | N/A
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | N | N/A | N/A
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | N | N/A | N/A
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | N | N/A | N/A
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | Y | ISO 14971 ISO 11607-1 ISO 11607-2 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or | N | N/A | N/A
— information on any known restrictions to combinations of devices and equipment; | N | N/A | N/A
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | N | N/A | N/A
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | N | N/A | N/A
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | N | N/A | N/A
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | Y | ISO 14971 ISO 10993-1 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | N | N/A | N/A
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | Y | ISO 14971 ISO 10993-1 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | N | N/A | N/A
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 11737-1 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
— physical hazards such as from sharps. | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 20417 | B010101 - Autologous blood collection bags B010102 - Homologous donor blood collection bags and kits B010201 - Whole blood, red blood cells or plasma transfer bags and kits B010202 - Platelets concentrate transfer bags and kits B0103 - Medullary blood collection, purification, cryopreservation bags and kits B010401 - Placental blood collection bags and kits B010402 - Placental blood purification bags and kits B010403 - Placental blood croypreservation bags and kits B0105 - Umbilical cord blood collection, purification, cryopreservation bags and kits
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | N | N/A | N/A
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A

## B03
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s) | Risk classification
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10555-1 ISO 10555-6 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 IEC 62366-1 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | B030101 - apheresis plasma collection devices | EU MDR, Class Ila
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 14971 | Medical devices — Application of risk management to medical devices | B030102 - Apheresis platelet collection devices | EU MDR, Class I / Class Ila
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10555-1 | Intravascular catheters — Sterile and single-use catheters Part 1: General requirements | B030103 - Apheresis leukocyte collection devices | N/A
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10555-6 | Intravascular catheters — Sterile and single-use catheters Part 6: Subcutaneous implanted ports | B030104 - Apheresis multiple blood component collection devices | N/A
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | B030201 - Plasmapheresis devices and kits | EU MDR, Class Ila / llb
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-3 | Biological evaluation of medical devices Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity | B030202 - Cytapheresis devices and kits | EU MDR, Class Ilb
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-4 | Biological evaluation of medical devices Part 4: Selection of tests for interactions with blood | B030203 - Single plasma components removal devices and kits | EU MDR, Class I / Ila / llb
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | N/A
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-11 | Biological evaluation of medical devices Part 11: Tests for systemic toxicity
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO 10993-18 | Biological evaluation of medical devices Part 18: Chemical characterization of medical device materials within a risk management process
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 10993-7 | Biological evaluation of medical devices Part 7: Ethylene oxide sterilization residuals
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 14937 | Sterilization of health care products — General requirements for characterization of a sterilizing agent and the development, validation and routine control of a sterilization process for medical devices
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | BS EN 556-1 | Sterilization of medical devices. Requirements for medical devices to be designated "STERILE" - Requirements for terminally sterilized medical devices
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 15223-1 | Medical devices — Symbols to be used with information to be supplied by the manufacturer Part 1: General requirements
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | ISO 20417 | Medical devices — Information to be supplied by the manufacturer
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | IEC 62304 | Medical device software — Software life cycle processes
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | IEC 60601-1 | Medical electrical equipment - Part 1: General requirements for basic safety and essential performance
|  |  |  | IEC 60601-1-2 | Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s)
10. Chemical, physical and biological properties
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to:
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10555-1 ISO 10555-6 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | N | N/A | N/A
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 10555-1 ISO 10555-6 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; | N | N/A | N/A
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 10555-1 ISO 10555-6 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(g) surface properties; and | Y | ISO 10555-1 ISO 10555-6 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(h) the confirmation that the device meets any defined chemical and/or physical specifications. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
10.4. Substances
10.4.1. Design and manufacture of devices
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that:
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 10555-1 ISO 10555-6 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | Y | ISO 14971 ISO 10555-1 ISO 10555-6 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | Y | ISO 14971 ISO 10555-1 ISO 10555-6 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 ISO 10555-1 ISO 10555-6 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | N | N/A | N/A
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | Does apheresis devices already having a specific microbial state?
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 ISO 11607-1 ISO 11607-2 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. |  | ISO 13485 ISO 14971 ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 ISO 11607-1 ISO 11607-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply: |  |  |  | Does the apheresis device comply with derivatives of tissues or cells of human origin, because it separates the human blood?
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; | N | N/A | N/A
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; | N | N/A | N/A
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. | N | N/A | N/A
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; | N | N/A | N/A
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; | N | N/A | N/A
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. | N | N/A | N/A
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. | N | N/A | N/A
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 10555-1 ISO 10555-6 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10555-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | Y | ISO 14971 IEC 62304 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-10 ISO 10993-11 ISO 10993-18 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | Y | ISO 14971 IEC 60601-1-2 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits | I'm unsure if it means the device needs to meet all requirements. I don't think the blood bag can meet calibration, but it adjusts the speed of the fluid.
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | Y | ISO 13485 ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | Y | ISO 13485 ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | N | N/A | N/A
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | N | N/A | N/A
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | N | N/A | N/A
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | N | N/A | N/A
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and | N | N/A | N/A
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | N | N/A | N/A
— ageing of the materials used, | N | N/A | N/A
— excess heat generated by the device, | N | N/A | N/A
— decreased accuracy of any measuring or control mechanism. | N | N/A | N/A
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | N | N/A | N/A
— the reliability of the source of energy. | N | N/A | N/A
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | N | N/A | N/A
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | N | N/A | N/A
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 62366-1 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. |  | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | N | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | N | N/A | N/A
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | N | N/A | N/A
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | N | N/A | N/A
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. | N | N/A | N/A
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— tissues or cells, or their derivatives, of human origin, or | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | N | N/A | N/A
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | N | N/A | N/A
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) a declaration that the device is in a sterile condition, | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) the method of sterilisation, | Y | ISO 10993-7 ISO 11135 ISO 14937 BS EN 556-1 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(d) the name and address of the manufacturer, | Y | ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(e) a description of the device, | Y | ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(g) if the device is custom-made, the words ‘custom-made device’, | Y | ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(h) the month and year of manufacture, | Y | ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(e) the performance characteristics of the device; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— identification of any consumable components and how to replace them, | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 | B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or | N | N/A | N/A
— information on any known restrictions to combinations of devices and equipment; | N | N/A | N/A
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | Y | IEC 60601-1 IEC 60601-1-2 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | Y | IEC 60601-1 IEC 60601-1-2 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | Y | ISO 14971 ISO 10993-1 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | N | N/A | N/A
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | Y | ISO 14971 ISO 10993-1 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | N | N/A | N/A
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
— physical hazards such as from sharps. | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | N | N/A | N/A
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | ISO 15223-1 ISO 20417 | B030101 - apheresis plasma collection devices B030102 - Apheresis platelet collection devices B030103 - Apheresis leukocyte collection devices B030104 - Apheresis multiple blood component collection devices B030201 - Plasmapheresis devices and kits B030202 - Cytapheresis devices and kits B030203 - Single plasma components removal devices and kits B030204 - Extracorporeal photochaemotherapy or photopheresis devices and kits

## J02
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s)
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 ISO 10555-1 ISO/TS 10974 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 IEC 62133-1 IEC 62133-2 IEC 62311 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | J020101 - Deep brain stimulation (DBS) implantable neurostimulators
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 14971 | Medical devices — Application of risk management to medical devices | J020102 - Deep brain stimulation (DBS) leads
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | J020201 - Radiofrequency implantable spinal neurostimulators
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | J020202 - Fully implantable spinal neurostimulators
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 10993-6 | Biological evaluation of medical devices Part 6: Tests for local effects after implantation | J020203 - Spinal neurostimulation leads
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization | J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 10993-17 | Biological evaluation of medical devices Part 17: Toxicological risk assessment of medical device constituents | J020302 - Vagal neurostimulation leads
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 10993-18 | Biological evaluation of medical devices Part 18: Chemical characterization of medical device materials within a risk management process | J020401 - Bladder incontinence neurostimulators
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 10993-23 | Biological evaluation of medical devices Part 23: Tests for irritation | J020402 - Intestinal incontinence neurostimulators
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 10555-1 | Intravascular catheters — Sterile and single-use catheters Part 1: General requirements | J020403 - Sacral spine neurostimulation leads
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO/TS 10974 | Assessment of the safety of magnetic resonance imaging for patients with an active implantable medical device | J0206 - Frenic nerve implantable neurostimulators
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 11070 | Sterile single-use intravascular introducers, dilators and guidewires | J020701 - Neurostimulators programmers
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 14708-1 | Implants for surgery — Active implantable medical devices Part 1: General requirements for safety, marking and for information to be provided by the manufacturer | J020801 - Inferior esophageal sphincter implantable neurostimulators
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 14708-3 | Implants for surgery — Active implantable medical devices Part 3: Implantable neurostimulators | J020802 - Gastric implantable obesity treatment neurostimulators
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 80369-6 | Small bore connectors for liquids and gases in healthcare applications Part 6: Connectors for neuraxial applications | J020803 - Gastric implantable nausea therapy neurostimulators
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | IEC 60529 | Degrees of protection provided by enclosures (IP Code) | J020804 - Gastric neurostimulation leads
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | IEC 60601-1 | Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | J0209 - Carotid sinus implantable neurostimulators and leads
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | IEC 60601-1-2 | Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests | J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 ISO 10555-1 ISO/TS 10974 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 IEC 62133-1 IEC 62133-2 IEC 62311 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | IEC 60601-1-6 | Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability | J0211 - Sleep apnea treatment implantable neurostimulators and leads
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 62133-1 IEC 62133-2 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | IEC 60601-1-11 | Medical electrical equipment - Part 1-11: General requirements for basic safety and essential performance - Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | IEC 60601-2-10 | Medical electrical equipment - Part 2-10: Particular requirements for the basic safety and essential performance of nerve and muscle stimulators
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | IEC 62133-1 | Secondary cells and batteries containing alkaline or other non-acid electrolytes - Safety requirements for portable sealed secondary cells, and for batteries made from them, for use in portable applications - Part 1: Nickel systems
|  |  |  | IEC 62133-2 | Secondary cells and batteries containing alkaline or other non-acid electrolytes - Safety requirements for portable sealed secondary cells, and for batteries made from them, for use in portable applications - Part 2: Lithium systems
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s) | IEC 62311 | Assessment of electronic and electrical equipment related to human exposure restrictions for electromagnetic fields (0 Hz - 300 GHz)
10. Chemical, physical and biological properties |  |  |  | IEC 62304 | Medical device software — Software life cycle processes
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to: |  |  |  | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 62133-1 IEC 62133-2 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 11137-1 | Sterilization of health care products — Radiation Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | Y | ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 11137-2 | Sterilization of health care products — Radiation Part 2: Establishing the sterilization dose
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 10555-1 ISO/TS 10974 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 62133-1 IEC 62133-2 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 10555-1 ISO/TS 10974 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 62133-1 IEC 62133-2 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 15223-1 | Medical devices — Symbols to be used with information to be supplied by the manufacturer Part 1: General requirements
(g) surface properties; and | Y | ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads | ISO 20417 | Medical devices — Information to be supplied by the manufacturer
(h) the confirmation that the device meets any defined chemical and/or physical specifications. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | ISO 7000 | Graphical symbols for use on equipment — Registered symbols
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
10.4. Substances
10.4.1. Design and manufacture of devices
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that:
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | N | N/A | N/A
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | N | N/A | N/A
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or | N | N/A | N/A
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. | N | N/A | N/A
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; | N | N/A | N/A
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; | N | N/A | N/A
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and | N | N/A | N/A
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. | N | N/A | N/A
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. | N | N/A | N/A
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. | N | N/A | N/A
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 62133-1 IEC 62133-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | N | N/A | N/A
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 11135 ISO 11137-1 ISO 11137-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. | N | N/A | N/A
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 11135 ISO 11137-1 ISO 11137-2 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 11135 ISO 11137-1 ISO 11137-2 ISO 11607-1 ISO 11607-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | N | N/A | N/A
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; | N | N/A | N/A
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; | N | N/A | N/A
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. | N | N/A | N/A
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; | N | N/A | N/A
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; | N | N/A | N/A
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. | N | N/A | N/A
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. | N | N/A | N/A
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | Y | ISO 14971 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 IEC 62133-1 IEC 62133-2 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 IEC 62311 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | Y | ISO 14971 IEC 62304 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | Y | ISO 14971 ISO/TS 10974 IEC 60601-1-2 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 IEC 60601-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | N | N/A | N/A
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. | Y | ISO 13485 ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. | Y | ISO 13485 ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. | Y | ISO 13485 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | Y | ISO 13485 ISO 14971 ISO/TS 10974 IEC 60601-1 IEC 60601-1-2 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | Y | ISO 14971 ISO/TS 10974 IEC 60601-1 IEC 60601-1-2 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | Y | ISO 14971 ISO/TS 10974 IEC 60601-1 IEC 60601-1-2 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | Y | ISO 13485 ISO 14971 ISO/TS 10974 IEC 60601-1 IEC 60601-1-2 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | Y | ISO 14971 IEC 62304 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | Y | ISO 14971 IEC 62304 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | ISO 14971 IEC 62304 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | Y | ISO 14971 IEC 60601-1-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | Y | ISO 14971 IEC 62304 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | Y | ISO 14971 IEC 60601-1-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | Y | ISO 14971 IEC 60601-1-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | Y | ISO 14971 IEC 60601-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | Y | ISO 13485 ISO 14971 IEC 62304 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | Y | ISO 14971 IEC 60529 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 IEC 62133-1 IEC 62133-2 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | Y | ISO 14971 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— ageing of the materials used, | Y | ISO 14971 ISO 10555-1 ISO 11070 ISO 14708-1 ISO 14708-3 ISO 80369-6 IEC 60529 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— excess heat generated by the device, | Y | ISO 14971 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— decreased accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62304 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | Y | ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-17 ISO 10993-18 ISO 10993-23 ISO/TS 10974 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— the reliability of the source of energy. | Y | ISO 14971 IEC 62133-1 IEC 62133-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | Y | ISO 14971 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | N | N/A | N/A
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 IEC 60601-1 IEC 62366-1 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020801 - Inferior esophageal sphincter implantable neurostimulators J020804 - Gastric neurostimulation leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. | N | N/A | N/A
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 60601-1-11 IEC 60601-2-10 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | Y | ISO 14971 IEC 62311 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | N | N/A | N/A
— tissues or cells, or their derivatives, of human origin, or | N | N/A | N/A
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | ISO 11135 ISO 11137-1 ISO 11137-2 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | N | N/A | N/A
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | N | N/A | N/A
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) a declaration that the device is in a sterile condition, | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(c) the method of sterilisation, | Y | ISO 11135 ISO 11137-1 ISO 11137-2 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(d) the name and address of the manufacturer, | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(e) a description of the device, | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(g) if the device is custom-made, the words ‘custom-made device’, | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(h) the month and year of manufacture, | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(e) the performance characteristics of the device; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | Y | IEC 62304 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | N | N/A | N/A
— identification of any consumable components and how to replace them, | N | N/A | N/A
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | N | N/A | N/A
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | N | N/A | N/A
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | N | N/A | N/A
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | N | N/A | N/A
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | N | N/A | N/A
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | Y | ISO 14971 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— information on any known restrictions to combinations of devices and equipment; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | Y | ISO/TS 10974 IEC 60601-1 IEC 60601-1-2 IEC 62311 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | Y | ISO/TS 10974 IEC 60601-1 IEC 60601-1-2 IEC 62311 ISO 11607-1 ISO 11607-2 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 IEC 62311 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | ISO 14971 IEC 62311 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | N | N/A | N/A
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | N | N/A | N/A
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | N | N/A | N/A
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
— physical hazards such as from sharps. | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | Y | ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | IEC 62304 ISO 15223-1 ISO 20417 ISO 7000 | J020101 - Deep brain stimulation (DBS) implantable neurostimulators J020102 - Deep brain stimulation (DBS) leads J020201 - Radiofrequency implantable spinal neurostimulators J020202 - Fully implantable spinal neurostimulators J020203 - Spinal neurostimulation leads J020301 - Drug-resistant epilepsy non-surgical therapy implantable neurostimulators J020302 - Vagal neurostimulation leads J020401 - Bladder incontinence neurostimulators J020402 - Intestinal incontinence neurostimulators J020403 - Sacral spine neurostimulation leads J0206 - Frenic nerve implantable neurostimulators J020701 - Neurostimulators programmers J020801 - Inferior esophageal sphincter implantable neurostimulators J020802 - Gastric implantable obesity treatment neurostimulators J020803 - Gastric implantable nausea therapy neurostimulators J020804 - Gastric neurostimulation leads J0209 - Carotid sinus implantable neurostimulators and leads J0210 - Peripheral nerve implantable neurostimulators and pain therapy leads J0211 - Sleep apnea treatment implantable neurostimulators and leads

## P09
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s)
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | P090103 - Glenoid components
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 14971 | Medical devices — Application of risk management to medical devices | P090104 - Shoulder prostheses humeral components
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | P090203 - Elbow prostheses humeral components
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10993-3 | Biological evaluation of medical devices Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity | P090204 - Elbow prostheses radial components
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | P090205 - Elbow prostheses ulnar components
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 10993-6 | Biological evaluation of medical devices Part 6: Tests for local effects after implantation | P090303 - Wrist prostheses carpal components
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization | P090304 - Wrist prostheses radial-ulnar components
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 10993-11 | Biological evaluation of medical devices Part 11: Tests for systemic toxicity | P090305 - Total wrist prostheses
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 10993-12 | Biological evaluation of medical devices Part 12: Sample preparation and reference materials | P090403 - Hand prostheses interphalange components
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 10993-18 | Biological evaluation of medical devices Part 18: Chemical characterization of medical device materials within a risk management process | P090404 - Hand prostheses trapezio-metacarpal components
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO 17664-1 | Processing of health care products — Information to be provided by the medical device manufacturer for the processing of medical devices Part 1: Critical and semi-critical medical devices | P090506 - Ankle prosthesis with mobile insert
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 21535 | Non-active surgical implants — Joint replacement implants — Specific requirements for hip-joint replacement implants | P090507 - Ankle prosthesis with fixed insert
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 5832-1 | Implants for surgery — Metallic materials Part 1: Wrought stainless steel | P090508 - Ankle prostheses spacers
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 5832-2 | Implants for surgery — Metallic materials Part 2: Unalloyed titanium | P090603 - Foot prostheses interphalangeal components
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 5832-3 | Implants for surgery — Metallic materials Part 3: Wrought titanium 6-aluminium 4-vanadium alloy | P090604 - Foot prostheses metatarsalphalangeal components
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | ISO 5832-9 | Implants for surgery — Metallic materials Part 9: Wrought high nitrogen stainless steel | P090605 - Foot endorthesis
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 6475 | Implants for surgery — Metal bone screws with asymmetrical thread and spherical under-surface — Mechanical requirements and test methods | P090701 - Spinal fusion systems
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 7206-4 | Implants for surgery Partial and total hip joint prostheses — Part 4: Determination of endurance properties and performance of stemmed femoral components | P090702 - Intervertebral discs replacement systems
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 7206-6 | Implants for surgery Partial and total hip joint prostheses — Part 6: Endurance properties testing and performance requirements of neck region of stemmed femoral components | P090703 - Implantable vertebral stabilisation or fixation systems
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices | P090803 - Hip prostheses acetabular components
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 11137-1 | Sterilization of health care products — Radiation Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices | P090804 - Hip prostheses femoral components
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | ISO 11137-2 | Sterilization of health care products — Radiation Part 2: Establishing the sterilization dose | P090805 - Hip prostheses spacers
|  |  |  | ISO 11137-3 | Sterilization of health care products — Radiation Part 3: Guidance on dosimetric aspects of development, validation and routine control | P090903 - Bicompartmental primary implant knee prostheses
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s) | ISO 11737-1 | Sterilization of health care products — Microbiological methods Part 1: Determination of a population of microorganisms on products | P090904 - Unicompartmental knee prostheses (unicondylar)
10. Chemical, physical and biological properties |  |  |  | ISO 11737-2 | Sterilization of health care products — Microbiological methods Part 2: Tests of sterility performed in the definition, validation and maintenance of a sterilization process | P090905 - Revision knee prostheses
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to: |  |  |  | ISO 17665 | Sterilization of health care products — Moist heat — Requirements for the development, validation and routine control of a sterilization process for medical devices | P090906 - Knee prostheses for large resections
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems | P090907 - Patello-femoral knee prostheses
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes | P090908 - Knee prostheses spacers
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | Y | ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use | ISO 20417 | Medical devices — Information to be supplied by the manufacturer | P0910 - Ligament prostheses
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use |  |  | P091201 - Osteosynthesis staples and anchors
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  |  |  | P091202 - Bone fixation nails
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use |  |  | P091203 - Bone fixation wires
(g) surface properties; and | Y | ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use |  |  | P091204 - External bone fixation systems
(h) the confirmation that the device meets any defined chemical and/or physical specifications. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  |  |  | P091205 - Bone fixation plates
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use |  |  | P091206 - Bone fixation screws, tendon and ligament synthesis screws
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use |  |  | P091301 - Orthopaedic implant reamers, single-use
10.4. Substances | [5 empty] | P091302 - Orthopaedic implant blades, single-use
10.4.1. Design and manufacture of devices | [5 empty] | P091303 - Orthopaedic implant drill bits, single-use
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that: | [5 empty] | P091304 - Cranial perforators, single-use
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use |  |  | P091305 - Bone saws, single-use
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | N | N/A | N/A
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | N | N/A | N/A
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or | N | N/A | N/A
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. | N | N/A | N/A
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; | N | N/A | N/A
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; | N | N/A | N/A
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and | N | N/A | N/A
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. | N | N/A | N/A
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. | N | N/A | N/A
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. | N | N/A | N/A
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | N | N/A | N/A
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 ISO 11737-1 ISO 11737-2 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 ISO 11737-1 ISO 11737-2 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 11137-1 ISO 11137-2 ISO 11137-3 ISO 17665 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. | N | N/A | N/A
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 11137-1 ISO 11137-2 ISO 11137-3 ISO 11737-1 ISO 11737-2 ISO 17665 ISO 11607-1 ISO 11607-2 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 11137-1 ISO 11137-2 ISO 11137-3 ISO 11737-1 ISO 11737-2 ISO 17665 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; | N | N/A | N/A
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; | N | N/A | N/A
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. | N | N/A | N/A
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; | N | N/A | N/A
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; | N | N/A | N/A
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. | N | N/A | N/A
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. | N | N/A | N/A
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | N | N/A | N/A
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | N | N/A | N/A
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | N | N/A | N/A
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. | Y | ISO 13485 ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | N | N/A | N/A
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | N | N/A | N/A
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | N | N/A | N/A
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | N | N/A | N/A
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | N | N/A | N/A
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | N | N/A | N/A
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | N | N/A | N/A
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— ageing of the materials used, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— excess heat generated by the device, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— decreased accuracy of any measuring or control mechanism. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | Y | ISO 10993-1 ISO 10993-3 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 10993-18 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
— the reliability of the source of energy. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | Y | ISO 14971 ISO 10993-1 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | Y | ISO 13485 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | N | N/A | N/A
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | Y | ISO 14971 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 ISO 17664-1 ISO 21535 ISO 5832-1 ISO 5832-2 ISO 5832-3 ISO 5832-9 ISO 6475 ISO 7206-4 ISO 7206-6 IEC 62366-1 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | N | N/A | N/A
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | N | N/A | N/A
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | N | N/A | N/A
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | N | N/A | N/A
— tissues or cells, or their derivatives, of human origin, or | N | N/A | N/A
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | ISO 11137-1 ISO 11137-2 ISO 11137-3 ISO 17665 ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | N | N/A | N/A
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | N | N/A | N/A
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(b) a declaration that the device is in a sterile condition, | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(c) the method of sterilisation, | Y | ISO 11137-1 ISO 11137-2 ISO 11137-3 ISO 17665 ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(d) the name and address of the manufacturer, | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(e) a description of the device, | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(g) if the device is custom-made, the words ‘custom-made device’, | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(h) the month and year of manufacture, | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(e) the performance characteristics of the device; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | N | N/A | N/A
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | N | N/A | N/A | Is the implant device applicable or not?
— identification of any consumable components and how to replace them, | N | N/A | N/A
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | N | N/A | N/A
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | N | N/A | N/A
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | N | N/A | N/A
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | N | N/A | N/A
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | N | N/A | N/A
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | Y | ISO 14971 ISO 11607-1 ISO 11607-2 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— information on any known restrictions to combinations of devices and equipment; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | N | N/A | N/A
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | N | N/A | N/A
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | N | N/A | N/A
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | N | N/A | N/A
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | N | N/A | N/A
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
— physical hazards such as from sharps. | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | Y | ISO 20417 | P090103 - Glenoid components P090104 - Shoulder prostheses humeral components P090203 - Elbow prostheses humeral components P090204 - Elbow prostheses radial components P090205 - Elbow prostheses ulnar components P090303 - Wrist prostheses carpal components P090304 - Wrist prostheses radial-ulnar components P090305 - Total wrist prostheses P090403 - Hand prostheses interphalange components P090404 - Hand prostheses trapezio-metacarpal components P090506 - Ankle prosthesis with mobile insert P090507 - Ankle prosthesis with fixed insert P090508 - Ankle prostheses spacers P090603 - Foot prostheses interphalangeal components P090604 - Foot prostheses metatarsalphalangeal components P090605 - Foot endorthesis P090701 - Spinal fusion systems P090702 - Intervertebral discs replacement systems P090703 - Implantable vertebral stabilisation or fixation systems P090803 - Hip prostheses acetabular components P090804 - Hip prostheses femoral components P090805 - Hip prostheses spacers P090903 - Bicompartmental primary implant knee prostheses P090904 - Unicompartmental knee prostheses (unicondylar) P090905 - Revision knee prostheses P090906 - Knee prostheses for large resections P090907 - Patello-femoral knee prostheses P090908 - Knee prostheses spacers P0910 - Ligament prostheses P091201 - Osteosynthesis staples and anchors P091202 - Bone fixation nails P091203 - Bone fixation wires P091204 - External bone fixation systems P091205 - Bone fixation plates P091206 - Bone fixation screws, tendon and ligament synthesis screws P091301 - Orthopaedic implant reamers, single-use P091302 - Orthopaedic implant blades, single-use P091303 - Orthopaedic implant drill bits, single-use P091304 - Cranial perforators, single-use P091305 - Bone saws, single-use
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).

## W02
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s)
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 EN 13532 EN 13612 ISO 20916 IEC 60601-1 IEC 60601-1-2 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61010-1 IEC 61010-2-101 IEC 61326-1 IEC 61326-2-6 IEC 62366-1 IEC 62479 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | W020101 - Chemistry instruments
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | ISO 14971 | Medical devices — Application of risk management to medical devices | W020102 - Immunochemistry instruments
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | W020103 - Integrated / Modular chemistry / Immunochemistry instruments
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | EN 13532 | General requirements for in vitro diagnostic medical devices for self-testing | W020104 - Blood gas analysis / Electrolyte instruments
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | EN 13612 | Performance evaluation of in vitro diagnostic medical devices | W020105 - Electrophoresis / Chromatography instruments
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | ISO 20916 | In vitro diagnostic medical devices — Clinical performance studies using specimens from human subjects — Good study practice | W020106 - Chemistry / Immunochemistry rapid tests + POC tests
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 60601-1 | Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | W020107 - Urinalysis (ur) - laboratory analysers
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 60601-1-2 | Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests | W020201 - Cell count
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 61000-4-2 | Electromagnetic compatibility (EMC) - Part 4-2: Testing and measurement techniques - Electrostatic discharge immunity test | W020202 - Hemostasis
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 61000-4-3 | Electromagnetic compatibility (EMC) - Part 4-3 : Testing and measurement techniques - Radiated, radio-frequency, electromagnetic field immunity test | W020203 - Blood grouping instruments
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | IEC 61000-4-6 | Electromagnetic compatibility (EMC) - Part 4-6: Testing and measurement techniques - Immunity to conducted disturbances, induced by radio-frequency fields | W020204 - Flow cytometry instruments
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 61000-4-8 | Electromagnetic compatibility (EMC) - Part 4-8: Testing and measurement techniques - Power frequency magnetic field immunity test | W020205 - Histology / Cytology
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 61010-1 | Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 1: General requirements | W020206 - Rapid test hematology / Histology / Cytology instruments
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 61010-2-101 | Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment | W020301 - Microbiology susceptibility / Identification
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 61326-1 | Electrical equipment for measurement, control and laboratory use - EMC requirements - Part 1: General requirements | W020302 - Blood cultures and mycobacteria
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | IEC 61326-2-6 | Electrical equipment for measurement, control and laboratory use - EMC requirements - Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment | W020303 - Mass - spectrometry system
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 62304 | Medical device software — Software life cycle processes | W020304 - Microbiology gram stainer
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices | W020305 - Plates streaker
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 EN 13532 EN 13612 IEC 60601-1 IEC 60601-1-2 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61010-1 IEC 61010-2-101 IEC 61326-1 IEC 61326-2-6 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 62479 | Assessment of the compliance of low-power electronic and electrical equipment with the basic restrictions related to human exposure to electromagnetic fields (10 MHz to 300 GHz) | W020306 - Urinary screening system
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 EN 13532 IEC 61010-1 IEC 61010-2-101 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 63000 | Amendment 1 - Technical documentation for the assessment of electrical and electronic products with respect to the restriction of hazardous substances | W0204 - Infectious immunology instruments
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | IEC 81001-5-1 | Health software and health IT systems safety, effectiveness and security Part 5-1: Security — Activities in the product life cycle | W020501 - Nucleic acid testing instruments except micro-arrays
|  |  |  | ISO 11737-1 | Sterilization of health care products — Microbiological methods Part 1: Determination of a population of microorganisms on products | W020502 - Micro-array instruments
|  |  |  | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices | W020503 - Nucleic acid testing instruments - rapid and point of care tests
Chapter II — Requirements regarding performance, design and manufacture | Apply Y/N | Relevant standard(s) | Relevant device(s) | ISO 11607-1 | Packaging for terminally sterilized medical devices Part 1: Requirements for materials, sterile barrier systems and packaging systems | W020601 - Robotic sample processing systems
9. Performance characteristics |  |  |  | ISO 11607-2 | Packaging for terminally sterilized medical devices Part 2: Validation requirements for forming, sealing and assembly processes | W020602 - Post-analysis sample archiving systems
9.1. Devices shall be designed and manufactured in such a way that they are suitable for the purposes referred to in point (2) of Article 2, as specified by the manufacturer, and suitable with regard to the performance they are intended to achieve, taking account of the generally acknowledged state of the art. They shall achieve the performances, as stated by the manufacturer and in particular, where applicable: |  |  |  | ISO 18113-1 | In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) Part 1: Terms, definitions, and general requirements | W020701 - Scales
(a) the analytical performance, such as, analytical sensitivity, analytical specificity, trueness (bias), precision (repeatability and reproducibility), accuracy (resulting from trueness and precision), limits of detection and quantitation, measuring range, linearity, cut-off, including determination of appropriate criteria for specimen collection and handling and control of known relevant endogenous and exogenous interference, cross-reactions; and | Y | EN 13532 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers | ISO 20417 | Medical devices — Information to be supplied by the manufacturer | W020702 - Cryogenics / Frozen
(b) the clinical performance, such as diagnostic sensitivity, diagnostic specificity, positive predictive value, negative predictive value, likelihood ratio, expected values in normal and affected populations. | Y | EN 13532 EN 13612 ISO 20916 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers |  |  | W020703 - Biological hoods and cabinets
9.2. The performance characteristics of the device shall be maintained during the lifetime of the device as indicated by the manufacturer. | Y | EN 13532 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers |  |  | W020704 - Shakers and homogenizers
9.3. Where the performance of devices depends on the use of calibrators and/or control materials, the metrological traceability of values assigned to calibrators and/or control materials shall be assured through suitable reference measurement procedures and/or suitable reference materials of a higher metrological order. Where available, metrological traceability of values assigned to calibrators and control materials shall be assured to certified reference materials or reference measurement procedures. | Y | EN 13612 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
9.4. The characteristics and performances of the device shall be specifically checked in the event that they may be affected when the device is used for the intended use under normal conditions:
(a) for devices for self-testing, performances obtained by laypersons; | Y | EN 13532 EN 13612 ISO 20916 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) for devices for near-patient testing, performances obtained in relevant environments (for example, patient home, emergency units, ambulances). | Y | EN 13532 EN 13612 ISO 20916 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
10. Chemical, physical and biological properties
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. | Y | EN 13532 EN 13612 IEC 60601-1 IEC 60601-1-2 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61010-1 IEC 61010-2-101 IEC 61326-1 IEC 61326-2-6 IEC 62366-1 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 18113-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
Particular attention shall be paid to the possibility of impairment of analytical performance due to physical and/or chemical incompatibility between the materials used and the specimens, analyte or marker to be detected (such as biological tissues, cells, body fluids and micro-organisms), taking account of the intended purpose of the device. | Y | ISO 14971 ISO 10993-1 EN 13532 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 EN 13532 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
10.3. Devices shall be designed and manufactured in such a way as to reduce to a level as low as reasonably practicable the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Special attention shall be given to substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, and to substances having endocrine disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council. | Y | ISO 14971 ISO 10993-1 EN 13532 IEC 63000 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
10.4. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device, taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 EN 13532 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
11. Infection and microbial contaminationA43:D54
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) allow easy and safe handling, | Y | ISO 14971 ISO 10993-1 EN 13532 EN 13612 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 14971 ISO 10993-1 EN 13532 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(c) prevent microbial contamination of the device during use and, in the case of specimen receptacles, the risk of contamination of the specimen. | Y | ISO 14971 ISO 10993-1 EN 13532 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
11.2. Devices labelled either as sterile or as having a specific microbial state shall be designed, manufactured and packaged to ensure that their sterile condition or microbial state is maintained under the transport and storage conditions specified by the manufacturer until that packaging is opened at the point of use, unless the packaging which maintains their sterile condition or microbial state is damaged. | Y | ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020202 - Hemostasis W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer
11.3. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | ISO 13485 ISO 14971 ISO 11737-1 ISO 11135 ISO 11607-1 ISO 11607-2 | W020202 - Hemostasis W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria
11.4. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | ISO 13485 ISO 14971 ISO 11737-1 ISO 11135 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
11.5. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
11.6. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
12. Devices incorporating materials of biological origin
Where devices include tissues, cells and substances of animal, human or microbial origin, the selection of sources, the processing, preservation, testing and handling of tissues, cells and substances of such origin and control procedures shall be carried out so as to provide safety for user or other person. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
In particular, safety with regard to microbial and other transmissible agents shall be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This might not apply to certain devices if the activity of the microbial and other transmissible agent are integral to the intended purpose of the device or when such elimination or inactivation process would compromise the performance of the device. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
13. Construction of devices and interaction with their environment
13.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. | Y | ISO 14971 IEC 62479 IEC 63000 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
13.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61010-1 IEC 61010-2-101 IEC 61326-1 IEC 61326-2-6 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 EN 13532 IEC 60601-1 IEC 60601-1-2 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61010-1 IEC 61010-2-101 IEC 61326-1 IEC 61326-2-6 IEC 62366-1 IEC 62479 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 EN 13532 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | Y | ISO 14971 IEC 62304 IEC 81001-5-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 EN 13532 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(f) the risk of incorrect identification of specimens and the risk of erroneous results due to, for example, confusing colour and/or numeric and/or character codings on specimen receptacles, removable parts and/or accessories used with devices in order to perform the test or assay as intended; | Y | ISO 14971 EN 13532 EN 13612 ISO 20916 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(g) the risks of any foreseeable interference with other devices. | Y | ISO 14971 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61326-1 IEC 61326-2-6 IEC 62479 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
13.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
13.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 EN 13532 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
13.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. | Y | ISO 13485 ISO 14971 EN 13532 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
13.6. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
13.7. The measuring, monitoring or display scale (including colour change and other visual indicators) shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. | Y | ISO 13485 ISO 14971 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
14. Devices with a measuring function
14.1. Devices having a primary analytical measuring function shall be designed and manufactured in such a way as to provide appropriate analytical performance in accordance with point (a) of Section 9.1 of Annex I, taking into account the intended purpose of the device. | Y | EN 13532 EN 13612 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
14.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. | Y | ISO 14971 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
15. Protection against radiation
15.1. Devices shall be designed, manufactured and packaged in such a way that exposure of users or other persons to radiation (intended, unintended, stray or scattered) is reduced as far as possible and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for diagnostic purposes. | Y | ISO 13485 ISO 14971 IEC 62479 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
15.2. When devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall as far as possible be:
(a) designed and manufactured in such a way as to ensure that the characteristics and the quantity of radiation emitted can be controlled and/or adjusted; and | Y | ISO 13485 ISO 14971 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61326-1 IEC 61326-2-6 IEC 62479 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
(b) fitted with visual displays and/or audible warnings of such emissions. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
15.3. The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | Y | ISO 14971 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61326-1 IEC 61326-2-6 IEC 62479 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
16. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
16.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | Y | ISO 14971 IEC 62304 IEC 81001-5-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
16.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | Y | ISO 14971 IEC 62304 IEC 81001-5-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
16.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
16.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | ISO 14971 IEC 62304 IEC 81001-5-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
17. Devices connected to or equipped with an energy source
17.1. For devices connected to or equipped with an energy source, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
17.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | Y | ISO 14971 EN 13532 IEC 62304 IEC 81001-5-1 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
17.3. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | Y | ISO 14971 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61326-1 IEC 61326-2-6 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
17.4. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | Y | ISO 14971 IEC 61000-4-2 IEC 61000-4-3 IEC 61000-4-6 IEC 61000-4-8 IEC 61326-1 IEC 61326-2-6 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
17.5. Devices shall be designed and manufactured in such a way as to avoid as far as possible the risk of accidental electric shocks to the user, or other person both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
18. Protection against mechanical and thermal risks
18.1. Devices shall be designed and manufactured in such a way as to protect users and other persons against mechanical risks. | Y | ISO 13485 ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
18.2. Devices shall be sufficiently stable under the foreseen operating conditions. They shall be suitable to withstand stresses inherent to the foreseen working environment, and to retain this resistance during the expected lifetime of the devices, subject to any inspection and maintenance requirements as indicated by the manufacturer. | Y | ISO 14971 EN 13532 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
18.3. Where there are risks due to the presence of moving parts, risks due to break-up or detachment, or leakage of substances, then appropriate protection means shall be incorporated. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
Any guards or other means included with the device to provide protection, in particular against moving parts, shall be secure and shall not interfere with access for the normal operation of the device, or restrict routine maintenance of the device as intended by the manufacturer. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
18.4. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
18.5. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
18.6. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
18.7. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 IEC 62366-1 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
18.8. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 IEC 61010-1 IEC 61010-2-101 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
19. Protection against the risks posed by devices intended for self-testing or near-patient testing
19.1. Devices intended for self-testing or near-patient testing shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to the intended user and the influence resulting from variation that can be reasonably anticipated in the intended user's technique and environment. The information and instructions provided by the manufacturer shall be easy for the intended user to understand and apply in order to correctly interpret the result provided by the device and to avoid misleading information. In the case of near-patient testing, the information and the instructions provided by the manufacturer shall make clear the level of training, qualifications and/or experience required by the user. | Y | ISO 14971 EN 13532 EN 13612 ISO 20916 IEC 62366-1 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
19.2. Devices intended for self-testing or near-patient testing shall be designed and manufactured in such a way as to:
(a) ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information; and | Y | ISO 14971 EN 13532 EN 13612 ISO 20916 IEC 62366-1 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
(b) reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | Y | ISO 14971 EN 13532 EN 13612 ISO 20916 IEC 62366-1 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
19.3. Devices intended for self-testing and near-patient testing shall, where feasible, include a procedure by which the intended user:
(a) can verify that, at the time of use, the device will perform as intended by the manufacturer, and | Y | ISO 14971 EN 13532 IEC 62366-1 ISO 18113-1 ISO 20417 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
(b) be warned if the device has failed to provide a valid result. | Y | ISO 14971 EN 13532 IEC 62366-1 ISO 18113-1 ISO 20417 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant standard(s) | Relevant device(s)
20. Label and instructions for use
20.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | EN 13532 IEC 62366-1 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit. If individual full labelling of each unit is not practicable, the information shall be set out on the packaging of multiple devices. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification or bar codes. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(d) Instructions for use shall be provided together with devices. However, in duly justified and exceptional cases instructions for use shall not be required or may be abbreviated if the device can be used safely and as intended by the manufacturer without any such instructions for use. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(e) Where multiple devices, with the exception of devices intended for self-testing or near-patient testing, are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) When the device is intended for professional use only, instructions for use may be provided to the user in non-paper format (e.g. electronic), except when the device is intended for near-patient testing. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(i) In the case of devices containing a substance or a mixture which may be considered as being dangerous, taking account of the nature and quantity of its constituents and the form under which they are present, relevant hazard pictograms and labelling requirements of Regulation (EC) No 1272/2008 shall apply. Where there is insufficient space to put all the information on the device itself or on its label, the relevant hazard pictograms shall be put on the label and the other information required by Regulation (EC) No 1272/2008 shall be given in the instructions for use. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(j) The provisions of Regulation (EC) No 1907/2006 on the safety data sheet shall apply, unless all relevant information, as appropriate, is already made available in the instructions for use. | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
20.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(e) an indication that the device is an in vitro diagnostic medical device, or if the device is a ‘device for performance study’, an indication of that fact; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(f) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(g) the UDI carrier referred to in Article 24 and Part C of Annex VI; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(h) an unambiguous indication of the time limit for using the device safely, without degradation of performance, expressed at least in terms of year and month and, where relevant, the day, in that order; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(i) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(j) where relevant, an indication of the net quantity of contents, expressed in terms of weight or volume, numerical count, or any combination of thereof, or other terms which accurately reflect the contents of the package; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(k) an indication of any special storage and/or handling condition that applies; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(l) where appropriate, an indication of the sterile state of the device and the sterilisation method, or a statement indicating any special microbial state or state of cleanliness; | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(n) if the instructions for use are not provided in paper form in accordance with point (f) of Section 20.1, a reference to their accessibility (or availability), and where applicable the website address where they can be consulted; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(o) where applicable, any particular operating instructions; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(p) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(q) if the device is intended for self-testing or near-patient testing, an indication of that fact; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(r) where rapid assays are not intended for self-testing or near-patient testing, the explicit exclusion hereof; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(s) where device kits include individual reagents and articles that are made available as separate devices, each of those devices shall comply with the labelling requirements contained in this Section and with the requirements of this Regulation; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(t) the devices and separate components shall be identified, where applicable in terms of batches, to allow all appropriate action to detect any potential risk posed by the devices and detachable components. As far as practicable and appropriate, the information shall be set out on the device itself and/or, where appropriate, on the sales packaging; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(u) the label for devices for self-testing shall bear the following particulars:
(i) the type of specimen(s) required to perform the test (e.g. blood, urine or saliva); | Y | EN 13532 ISO 18113-1 ISO 20417 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
(ii) the need for additional materials for the test to function properly; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
(iii) contact details for further advice and assistance. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
The name of devices for self-testing shall not reflect an intended purpose other than that specified by the manufacturer. | Y | EN 13532 ISO 18113-1 ISO 20417 | W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020503 - Nucleic acid testing instruments - rapid and point of care tests
20.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(b) a declaration that the device is in a sterile condition, | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(c) the method of sterilisation, | Y | EN 13532 ISO 11737-1 ISO 11135 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(d) the name and address of the manufacturer, | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(e) a description of the device, | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(f) the month and year of manufacture, | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(g) an unambiguous indication of the time limit for using the device safely, expressed at least in terms of year and month and, where relevant, the day, in that order, | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
20.4. Information in the instructions for use
20.4.1. The instructions for use shall contain all of the following particulars:
(a) the name or trade name of the device; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) the details strictly necessary for the user to uniquely identify the device; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(c) the device's intended purpose:
(i) what is detected and/or measured; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ii) its function (e.g. screening, monitoring, diagnosis or aid to diagnosis, prognosis, prediction, companion diagnostic); | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(iii) the specific information that is intended to be provided in the context of:
— a physiological or pathological state; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
— congenital physical or mental impairments; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— the predisposition to a medical condition or a disease; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— the determination of the safety and compatibility with potential recipients; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— the prediction of treatment response or reactions; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— the definition or monitoring of therapeutic measures; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(iv) whether it is automated or not; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(v) whether it is qualitative, semi-quantitative or quantitative; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(vi) the type of specimen(s) required; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(vii) where applicable, the testing population; and | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(viii) for companion diagnostics, the International Non-proprietary Name (INN) of the associated medicinal product for which it is a companion test. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(d) an indication that the device is an in vitro diagnostic medical device, or, if the device is a ‘device for performance study’, an indication of that fact; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(e) the intended user, as appropriate (e.g. self-testing, near patient and laboratory professional use, healthcare professionals); | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(f) the test principle; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(g) a description of the calibrators and controls and any limitation upon their use (e.g. suitable for a dedicated instrument only); | N | N/A | N/A
(h) a description of the reagents and any limitation upon their use (e.g. suitable for a dedicated instrument only) and the composition of the reagent product by nature and amount or concentration of the active ingredient(s) of the reagent(s) or kit as well as a statement, where appropriate, that the device contains other ingredients which might influence the measurement; | N | N/A | N/A
(i) a list of materials provided and a list of special materials required but not provided; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(j) for devices intended for use in combination with or installed with or connected to other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a validated and safe combination, including key performance characteristics, and/or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— information on any known restrictions to combinations of devices and equipment. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(k) an indication of any special storage (e.g. temperature, light, humidity, etc.) and/or handling conditions which apply; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(l) in-use stability which may include the storage conditions, and shelf life following the first opening of the primary container, together with the storage conditions and stability of working solutions, where this is relevant; | Y | EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(m) if the device is supplied as sterile, an indication of its sterile state, the sterilisation method and instructions in the event of the sterile packaging being damaged before use; | Y | EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020205 - Histology / Cytology W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020304 - Microbiology gram stainer W020305 - Plates streaker
(n) information that allows the user to be informed of any warnings, precautions, measures to be taken and limitations of use regarding the device. That information shall cover, where appropriate:
(i) warnings, precautions and/or measures to be taken in the event of malfunction of the device or its degradation as suggested by changes in its appearance that may affect performance, | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ii) warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(iii) warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(iv) precautions related to materials incorporated into the device that contain or consist of CMR substances, or endocrine disrupting substances or that could result in sensitisation or an allergic reaction by the patient or user, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
(v) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union, |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(vi) if the device is reusable, information on the appropriate processes to allow reuse, including cleaning, disinfection, decontamination, packaging and, where appropriate, the validated method of re-sterilisation. Information shall be provided to identify when the device should no longer be reused, such as signs of material degradation or the maximum number of allowable reuses; | Y | ISO 14971 EN 13532 ISO 11607-1 ISO 11607-2 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(o) any warnings and/or precautions related to potentially infectious material that is included in the device; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(p) where relevant, requirements for special facilities, such as a clean room environment, or special training, such as on radiation safety, or particular qualifications of the intended user; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(q) conditions for collection, handling, and preparation of the specimen; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(r) details of any preparatory treatment or handling of the device before it is ready for use, such as sterilisation, final assembly, calibration, etc., for the device to be used as intended by the manufacturer; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(s) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, including cleaning and disinfection; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
— identification of any consumable components and how to replace them; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
— methods for mitigating the risks encountered by persons involved in installing, calibrating or servicing devices. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(t) where applicable, recommendations for quality control procedures; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(u) the metrological traceability of values assigned to calibrators and control materials, including identification of applied reference materials and/or reference measurement procedures of higher order and information regarding maximum (self-allowed) batch to batch variation provided with relevant figures and units of measure; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(v) assay procedure including calculations and interpretation of results and where relevant if any confirmatory testing shall be considered; where applicable, the instructions for use shall be accompanied by information regarding batch to batch variation provided with relevant figures and units of measure; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(w) analytical performance characteristics, such as analytical sensitivity, analytical specificity, trueness (bias), precision (repeatability and reproducibility), accuracy (resulting from trueness and precision), limits of detection and measurement range, (information needed for the control of known relevant interferences, cross-reactions and limitations of the method), measuring range, linearity and information about the use of available reference measurement procedures and materials by the user; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(x) clinical performance characteristics as defined in Section 9.1 of this Annex; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(y) the mathematical approach upon which the calculation of the analytical result is made; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(z) where relevant, clinical performance characteristics, such as threshold value, diagnostic sensitivity and diagnostic specificity, positive and negative predictive value; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(aa) where relevant, reference intervals in normal and affected populations; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ab) information on interfering substances or limitations (e.g. visual evidence of hyperlipidaemia or haemolysis, age of specimen) that may affect the performance of the device; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ac) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories, and the consumables used with it, if any. This information shall cover, where appropriate:
(i) infection or microbial hazards, such as consumables contaminated with potentially infectious substances of human origin; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ii) environmental hazards such as batteries or materials that emit potentially hazardous levels of radiation); | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(iii) physical hazards such as explosion. | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ad) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business at which he can be contacted and its location be established, together with a telephone number and/or fax number and/or website address to obtain technical assistance; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ae) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use, with a clear indication of the introduced modifications; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(af) a notice to the user that any serious incident that has occurred in relation to the device shall be reported to the manufacturer and the competent authority of the Member State in which the user and/or the patient is established; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ag) where device kits include individual reagents and articles that may be made available as separate devices, each of these devices shall comply with the instructions for use requirements contained in this Section and with the requirements of this Regulation; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(ah) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | ISO 14971 EN 13532 IEC 62304 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020206 - Rapid test hematology / Histology / Cytology instruments W020303 - Mass - spectrometry system W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems
20.4.2. In addition, the instructions for use for devices intended for self-testing shall comply with all of the following principles:
(a) details of the test procedure shall be given, including any reagent preparation, specimen collection and/or preparation and information on how to run the test and interpret the results; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(b) specific particulars may be omitted provided that the other information supplied by the manufacturer is sufficient to enable the user to use the device and to understand the result(s) produced by the device; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) the device's intended purpose shall provide sufficient information to enable the user to understand the medical context and to allow the intended user to make a correct interpretation of the results; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(d) the results shall be expressed and presented in a way that is readily understood by the intended user; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(e) information shall be provided with advice to the user on action to be taken (in case of positive, negative or indeterminate result), on the test limitations and on the possibility of false positive or false negative result. Information shall also be provided as to any factors that can affect the test result such as age, gender, menstruation, infection, exercise, fasting, diet or medication; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(f) the information provided shall include a statement clearly directing that the user should not take any decision of medical relevance without first consulting the appropriate healthcare professional, information on disease effects and prevalence, and, where available, information specific to the Member State(s) where the device is placed on the market on where a user can obtain further advice such as national helplines, websites; | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers
(g) for devices intended for self-testing used for the monitoring of a previously diagnosed existing disease or condition, the information shall specify that the patient should only adapt the treatment if he has received the appropriate training to do so. | Y | ISO 14971 EN 13532 ISO 18113-1 ISO 20417 | W020101 - Chemistry instruments W020102 - Immunochemistry instruments W020103 - Integrated / Modular chemistry / Immunochemistry instruments W020104 - Blood gas analysis / Electrolyte instruments W020105 - Electrophoresis / Chromatography instruments W020106 - Chemistry / Immunochemistry rapid tests + POC tests W020107 - Urinalysis (ur) - laboratory analysers W020201 - Cell count W020202 - Hemostasis W020203 - Blood grouping instruments W020204 - Flow cytometry instruments W020205 - Histology / Cytology W020206 - Rapid test hematology / Histology / Cytology instruments W020301 - Microbiology susceptibility / Identification W020302 - Blood cultures and mycobacteria W020303 - Mass - spectrometry system W020304 - Microbiology gram stainer W020305 - Plates streaker W020306 - Urinary screening system W0204 - Infectious immunology instruments W020501 - Nucleic acid testing instruments except micro-arrays W020502 - Micro-array instruments W020503 - Nucleic acid testing instruments - rapid and point of care tests W020601 - Robotic sample processing systems W020602 - Post-analysis sample archiving systems W020701 - Scales W020702 - Cryogenics / Frozen W020703 - Biological hoods and cabinets W020704 - Shakers and homogenizers

## Z11
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standard(s) | Name | Medical device(s)
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y | ISO 13485 ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-10 ISO 10993-23 ISO/IEC 15444-1 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-3 IEC 60601-1-6 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC TS 60601-4-2 IEC 60825-1 IEC 60976 IEC 61217 IEC 61223-3-4 IEC 62083 IEC 62133-2 IEC 62274 IEC 62359 IEC 62366-1 IEC 62464-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | Z110101 - Linear accelerators
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | ISO 14971 | Medical devices — Application of risk management to medical devices | Z110102 - Radiotherapy simulation instruments
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | Z110103 - Brachytherapy radiation instruments
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | Z110104 - Radiotherapy treatment planning instruments
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization | Z110105 - Intravascular radiotherapy instruments
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | ISO 10993-23 | Biological evaluation of medical devices Part 23: Tests for irritation | Z110107 - Radiosurgery instruments
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | ISO/IEC 15444-1 | Information technology — JPEG 2000 image coding system Part 1: Core coding system | Z110108 - Tomotherapy instruments
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-1 | Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Z110109 - Magnetic resonance radiotherapy instruments
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-1-2 | Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests | Z110110 - Instruments for proton therapy
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-1-3 | Medical electrical equipment - Part 1-3: General requirements for basic safety and essential performance - Collateral Standard: Radiation protection in diagnostic X-ray equipment | Z110201 - Computerised gamma cameras
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | IEC 60601-1-6 | Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability | Z110202 - Integrated ct/gamma camera systems
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-1 | Medical electrical equipment - Part 2-1: Particular requirements for the basic safety and essential performance of electron accelerators in the range 1 MeV to 50 MeV | Z110203 - Integrated ct/pet systems
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-8 | Medical electrical equipment - Part 2-8: Particular requirements forthe basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV | Z110204 - Positron emission tomography instruments
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-28 | Medical electrical equipment - Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis | Z110205 - Radioisotope scanners
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-33 | Medical electrical equipment - Part 2-33: Particular requirements for the basic safety and essential performance of magnetic resonance equipment for medical diagnosis | Z110206 - Intraoperative radioisotope scanners
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | IEC 60601-2-37 | Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment | Z110207 - Integrated mri/pet systems
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-43 | Medical electrical equipment - Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures | Z110301 - Digital angiography systems
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-44 | Medical electrical equipment - Part 2-44: Particular requirements for the basic safety and essential performance of X-ray equipment for computed tomography | Z110302 - Mammography systems
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-3 IEC 60601-1-6 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC TS 60601-4-2 IEC 60825-1 IEC 60976 IEC 62133-2 IEC 62359 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-45 | Medical electrical equipment - Part 2-45: Particular requirements for the basic safety and essential performance of mammographic X-ray equipment and mammographic stereotactic devices | Z110303 - Instruments for orthopantomography and panoramic dental radiology
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 13485 ISO 14971 IEC 60601-1-3 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-54 | Medical electrical equipment - Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy | Z110304 - Instruments for endoral radiology
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 60601-2-63 | Medical electrical equipment - Part 2-63: Particular requirements for the basic safety and essential performance of dental extra-oral X-ray equipment | Z110305 - Radiology instruments
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s). |  | IEC 60601-2-65 | Medical electrical equipment - Part 2-65: Particular requirements for the basic safety and essential performance of dental intra-oral X-ray equipment | Z110306 - Computed tomographs (CT)
|  |  |  | IEC 60601-2-68 | Medical electrical equipment - Part 2-68: Particular requirements for the basic safety and essential performance of X-ray-based image-guided radiotherapy equipment for use with electron accelerators, light ion beam therapy equipment and radionuclide beam therapy equipment | Z110307 - Remote control tables
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s) | IEC TS 60601-4-2 | Medical electrical equipment - Part 4-2: Guidance and interpretation - Electromagnetic immunity: performance of medical electrical equipment and medical electrical systems | Z110310 - Multifunctional systems for traditional radiology
10. Chemical, physical and biological properties |  |  |  | IEC 60825-1 | Safety of laser products - Part 1: Equipment classification and requirements | Z110311 - Direct digital radiology (DR) systems
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to: |  |  |  | IEC 60976 | Medical electrical equipment - Medical electron accelerators - Functional performance characteristics | Z110401 - Ultrasound scanners
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-10 ISO 10993-23 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 61217 | Radiotherapy equipment - Coordinates, movements and scales | Z110402 - Ultrasound probes
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-10 ISO 10993-23 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 61223-3-4 | Evaluation and routine testing in medical imaging departments - Part 3-4: Acceptance tests - Imaging performance of dental X-ray equipment | Z110501 - Magnetic resonance (MR) systems
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | N | N/A | N/A | IEC 62083 | Medical electrical equipment - Requirements for the safety of radiotherapy treatment planning systems | Z110601 - Laser imaging scanner
(d) the impact of processes on material properties; | Y | ISO 13485 IEC 61217 IEC 62083 IEC 62133-2 IEC 62274 IEC 62359 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 62133-2 | Secondary cells and batteries containing alkaline or other non-acid electrolytes - Safety requirements for portable sealed secondary cells, and for batteries made from them, for use in portable applications - Part 2: Lithium systems | Z110602 - Digital computed radiography (CR) systems
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; | Y | IEC 61217 IEC 62083 IEC 62274 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110703 - X–ray film viewer Z110705 - Laser imagers for bioimaging Z119011 - Mammography ecography or radiography stereotactic systems | IEC 62274 | Medical electrical equipment - Safety of radiotherapy record and verify systems | Z110603 - Picture archiving and communication systems
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | IEC 61217 IEC 62083 IEC 62133-2 IEC 62274 IEC 62359 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 62304 | Medical device software — Software life cycle processes | Z110701 - Automatic film loading systems
(g) surface properties; and | N | N/A | N/A | IEC 62311 | Assessment of electronic and electrical equipment related to human exposure restrictions for electromagnetic fields (0 Hz - 300 GHz) | Z110702 - Diaphanoscope
(h) the confirmation that the device meets any defined chemical and/or physical specifications. | Y | IEC 61217 IEC 62083 IEC 62274 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110703 - X–ray film viewer Z110705 - Laser imagers for bioimaging Z119011 - Mammography ecography or radiography stereotactic systems | IEC 62359 | Ultrasonics - Field characterization - Test methods for the determination of thermal and mechanical indices related to medical diagnostic ultrasonic fields | Z110703 - X–ray film viewer
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 ISO 10993-1 ISO 10993-5 ISO 10993-10 ISO 10993-23 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices | Z110704 - X–ray film players
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages | IEC 62464-1 | Magnetic resonance equipment for medical imaging - Part 1: Determination of essential image quality parameters | Z110705 - Laser imagers for bioimaging
10.4. Substances |  |  |  | IEC 82304-1 | Health software Part 1: General requirements for product safety | Z110706 - Video or digital bioimaging players
10.4.1. Design and manufacture of devices |  |  |  | ISO 15223-1 | Medical devices — Symbols to be used with information to be supplied by the manufacturer Part 1: General requirements | Z110707 - Processors
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that: |  |  |  | ISO 20417 | Medical devices — Information to be supplied by the manufacturer | Z119001 - Bone densitometry systems
— are invasive and come into direct contact with the human body, | N | N/A | N/A |  |  | Z119002 - Dosimeters
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | N | N/A | N/A |  |  | Z119003 - Bioimaging photography instruments
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | N | N/A | N/A |  |  | Z119004 - Microfilm image readers
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2: | [5 empty] | Z119005 - Personal dosimeter readers
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s). |  |  |  | Z119007 - Image acquisition modules
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s). |  |  |  | Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances | [5 empty] | Z119010 - Laser systems for patient positioning
The justification for the presence of such substances shall be based upon: | [5 empty] | Z119011 - Mammography ecography or radiography stereotactic systems
(a) an analysis and estimation of potential patient or user exposure to the substance; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s). |  |  |  | Z119012 - Closed circuit television systems
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s). |  |  |  | Z119013 - Thermographs
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s). |  |  |  | Z119014 - Video recorders for bioimages
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the plastic additives used in the medical device(s).
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 13485 ISO 14971 IEC 62083 IEC 62274 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | N | N/A | N/A
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | N | N/A | N/A
(b) allow easy and safe handling, | Y | ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 10993-1 ISO 10993-5 ISO 10993-10 ISO 10993-23 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 10993-1 ISO 10993-5 ISO 10993-10 ISO 10993-23 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. | N | N/A | N/A
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | N | N/A | N/A
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | N | N/A | N/A
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | N | N/A | N/A
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | Y | ISO 13485 ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; | N | N/A | N/A
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; | N | N/A | N/A
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. | N | N/A | N/A
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; | N | N/A | N/A
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; | N | N/A | N/A
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. | N | N/A | N/A
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. | N | N/A | N/A
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | Y | ISO 14971 IEC 62311 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-3 IEC 60601-1-6 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC TS 60601-4-2 IEC 60825-1 IEC 60976 IEC 62133-2 IEC 62359 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110703 - X–ray film viewer Z110705 - Laser imagers for bioimaging Z119011 - Mammography ecography or radiography stereotactic systems
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-3 IEC 60601-1-6 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC TS 60601-4-2 IEC 60825-1 IEC 60976 IEC 62311 IEC 62359 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 ISO 10993-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | Y | ISO 14971 ISO/IEC 15444-1 IEC 62304 IEC 82304-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10993-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | Y | ISO 14971 IEC 60601-1-2 IEC TS 60601-4-2 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. | Y | ISO 13485 ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. | Y | ISO 13485 ISO 14971 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. | Y | IEC 60601-1-3 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 IEC 62464-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. | Y | ISO 14971 IEC 62304 IEC 62311 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | Y | ISO 13485 ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC TS 60601-4-2 IEC 62311 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC TS 60601-4-2 IEC 62083 IEC 62274 IEC 62311 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | Y | ISO 14971 IEC 60601-1-3 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 IEC 61223-3-4 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | Y | ISO 14971 IEC 62083 IEC 62274 IEC 62311 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | Y | ISO 13485 ISO 14971 IEC 62083 IEC 62274 IEC 62311 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | Y | ISO 14971 IEC 60601-1-3 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 IEC 61223-3-4 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | Y | ISO 14971 IEC 60601-1-3 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 IEC 61223-3-4 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | Y | ISO 14971 IEC 60601-1-3 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 IEC 61223-3-4 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | Y | ISO 14971 IEC 60601-1-3 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC 60825-1 IEC 60976 IEC 61223-3-4 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | Y | ISO 14971 ISO/IEC 15444-1 IEC 62304 IEC 82304-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | Y | ISO 14971 ISO/IEC 15444-1 IEC 62304 IEC 82304-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | Y | ISO 14971 ISO/IEC 15444-1 IEC 62304 IEC 82304-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | ISO 14971 ISO/IEC 15444-1 IEC 62304 IEC 82304-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | N | N/A | N/A
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and | N | N/A | N/A
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | N | N/A | N/A
— ageing of the materials used, | N | N/A | N/A
— excess heat generated by the device, | N | N/A | N/A
— decreased accuracy of any measuring or control mechanism. | N | N/A | N/A
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | N | N/A | N/A
— the reliability of the source of energy. | N | N/A | N/A
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | N | N/A | N/A
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | N | N/A | N/A
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 13485 ISO 14971 IEC 61217 IEC 62359 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. | Y | ISO 13485 ISO 14971 IEC 61217 IEC 62083 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | Y | ISO 13485 ISO 14971 IEC 62083 IEC 62366-1 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC 60601-1-3 IEC 60601-2-1 IEC 60601-2-8 IEC 60601-2-28 IEC 60601-2-33 IEC 60601-2-37 IEC 60601-2-43 IEC 60601-2-44 IEC 60601-2-45 IEC 60601-2-54 IEC 60601-2-63 IEC 60601-2-65 IEC 60601-2-68 IEC TS 60601-4-2 IEC 60825-1 IEC 60976 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. | Y | ISO 14971 IEC 61217 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 IEC 60601-1 IEC 60601-1-2 IEC TS 60601-4-2 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | N | N/A | N/A
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | N | N/A | N/A
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | N | N/A | N/A
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | IEC 62366-1 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | N | N/A | N/A
— tissues or cells, or their derivatives, of human origin, or | N | N/A | N/A
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(k) an indication of any special storage and/or handling condition that applies; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | N | N/A | N/A
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | N | N/A | N/A
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | N | N/A | N/A
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | N | N/A | N/A
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | N | N/A | N/A
(b) a declaration that the device is in a sterile condition, | N | N/A | N/A
(c) the method of sterilisation, | N | N/A | N/A
(d) the name and address of the manufacturer, | N | N/A | N/A
(e) a description of the device, | N | N/A | N/A
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | N | N/A | N/A
(g) if the device is custom-made, the words ‘custom-made device’, | N | N/A | N/A
(h) the month and year of manufacture, | N | N/A | N/A
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | N | N/A | N/A
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | N | N/A | N/A
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(c) where applicable, a specification of the clinical benefits to be expected. | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(e) the performance characteristics of the device; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | Y | ISO/IEC 15444-1 IEC 62304 IEC 82304-1 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— identification of any consumable components and how to replace them, | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | N | N/A | N/A
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | N | N/A | N/A
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— information on any known restrictions to combinations of devices and equipment; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | Y | IEC 60601-1 IEC 60601-1-2 IEC TS 60601-4-2 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | Y | IEC 60601-1 IEC 60601-1-2 IEC TS 60601-4-2 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | ISO 14971 IEC 62311 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | ISO 14971 IEC 62311 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | N | N/A | N/A
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | N | N/A | N/A
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | N | N/A | N/A
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | Y | IEC 62304 IEC 62311 IEC 62359 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems
— physical hazards such as from sharps. | N | N/A | N/A
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | ISO 14971 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | N | N/A | N/A
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | ISO/IEC 15444-1 IEC 62304 IEC 82304-1 ISO 15223-1 ISO 20417 | Z110101 - Linear accelerators Z110102 - Radiotherapy simulation instruments Z110103 - Brachytherapy radiation instruments Z110104 - Radiotherapy treatment planning instruments Z110105 - Intravascular radiotherapy instruments Z110107 - Radiosurgery instruments Z110108 - Tomotherapy instruments Z110109 - Magnetic resonance radiotherapy instruments Z110110 - Instruments for proton therapy Z110201 - Computerised gamma cameras Z110202 - Integrated ct/gamma camera systems Z110203 - Integrated ct/pet systems Z110204 - Positron emission tomography instruments Z110205 - Radioisotope scanners Z110206 - Intraoperative radioisotope scanners Z110207 - Integrated mri/pet systems Z110301 - Digital angiography systems Z110302 - Mammography systems Z110303 - Instruments for orthopantomography and panoramic dental radiology Z110304 - Instruments for endoral radiology Z110305 - Radiology instruments Z110306 - Computed tomographs (CT) Z110307 - Remote control tables Z110310 - Multifunctional systems for traditional radiology Z110311 - Direct digital radiology (DR) systems Z110401 - Ultrasound scanners Z110402 - Ultrasound probes Z110501 - Magnetic resonance (MR) systems Z110601 - Laser imaging scanner Z110602 - Digital computed radiography (CR) systems Z110603 - Picture archiving and communication systems Z110701 - Automatic film loading systems Z110702 - Diaphanoscope Z110703 - X–ray film viewer Z110704 - X–ray film players Z110705 - Laser imagers for bioimaging Z110706 - Video or digital bioimaging players Z110707 - Processors Z119001 - Bone densitometry systems Z119002 - Dosimeters Z119003 - Bioimaging photography instruments Z119004 - Microfilm image readers Z119005 - Personal dosimeter readers Z119007 - Image acquisition modules Z119008 - Bioimaging visualization and/or diagnostic report viewing monitors Z119010 - Laser systems for patient positioning Z119011 - Mammography ecography or radiography stereotactic systems Z119012 - Closed circuit television systems Z119013 - Thermographs Z119014 - Video recorders for bioimages

## C01- Arterio
Annex I — General safety and performance requirements (GSPR) |  |  |  | Standards | N/A | Device(s)
Chapter I — General requirements | Apply Y/N | Relevant Standard(s) | Relevant device(s) | Standards | Name | Medical device(s) | Risk classification
1. Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | Y |  | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 13485 | Medical devices — Quality management systems — Requirements for regulatory purposes | C010101 - Peripheral I.V. catheters
2. The requirement in this Annex to reduce risks as far as possible means the reduction of risks as far as possible without adversely affecting the benefit-risk ratio. | Y |  | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 14971 | Medical devices — Application of risk management to medical devices | C010101 - Peripheral I.V. catheters
3. Manufacturers shall establish, implement, document and maintain a risk management system. |  |  |  | ISO 10555-1 | Intravascular catheters — Sterile and single-use catheters Part 1: General requirements | C010103 - Integrated peripheral I.V. devices
Risk management shall be understood as a continuous iterative process throughout the entire lifecycle of a device, requiring regular systematic updating. In carrying out risk management manufacturers shall: |  |  |  | ISO 10555-5 | Intravascular catheters — Sterile and single-use catheters Part 5: Over-needle peripheral catheters | C010104 - Umbilical I.V. cannulas
(a) establish and document a risk management plan for each device; | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-1 | Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | C010201 - Central I.V. catheters, peripheral access
(b) identify and analyse the known and foreseeable hazards associated with each device; | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-3 | Biological evaluation of medical devices Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity | C010202 - Central I.V. catheters, non-tunneled
(c) estimate and evaluate the risks associated with, and occurring during, the intended use and during reasonably foreseeable misuse; | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-4 | Biological evaluation of medical devices Part 4: Selection of tests for interactions with blood | C010203 - Central venous catheters, partially tunneled
(d) eliminate or control the risks referred to in point (c) in accordance with the requirements of Section 4; | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-5 | Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity | C010204 - Subcutaneous implantable venous access port systems
(e)evaluate the impact of information from the production phase and, in particular, from the post-market surveillance system, on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability; and | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-6 | Biological evaluation of medical devices Part 6: Tests for local effects after implantation | C010301 - Radial cannulas
(f) based on the evaluation of the impact of the information referred to in point (e), if necessary amend control measures in line with the requirements of Section 4. | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-10 | Biological evaluation of medical devices Part 10: Tests for skin sensitization | C010302 - Umbilical cannulas
4. Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: |  |  |  | ISO 10993-11 | Biological evaluation of medical devices Part 11: Tests for systemic toxicity | C010303 - Femoral cannulas
(a) eliminate or reduce risks as far as possible through safe design and manufacture; | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 10993-12 | Biological evaluation of medical devices Part 12: Sample preparation and reference materials | C010401 - Cardiac angiography devices
(b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 14937 | Sterilization of health care products — General requirements for characterization of a sterilizing agent and the development, validation and routine control of a sterilization process for medical devices | C010402 - Peripheral angiography devices
(c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 23908 | Sharps injury protection — Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling | C010403 - Fractional flow reserve (ffr) measurement devices
Manufacturers shall inform users of any residual risks. | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 7864 | Sterile hypodermic needles for single use — Requirements and test methods | C010501 - Vena cava filters
5. In eliminating or reducing risks related to use error, the manufacturer shall: |  |  |  | ISO 80369-7 | Small-bore connectors for liquids and gases in healthcare applications Part 7: Connectors for intravascular or hypodermic applications | C010502 - Intravascular emobilic protection catheters and systems
(a) reduce as far as possible the risks related to the ergonomic features of the device and the environment in which the device is intended to be used (design for patient safety), and | Y | ISO 14971 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 80369-20 | Small-bore connectors for liquids and gases in healthcare applications Part 20: Common test methods
(b) give consideration to the technical knowledge, experience, education, training and use environment, where applicable, and the medical and physical conditions of intended users (design for lay, professional, disabled or other users). | Y | ISO 14971 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 8536-4 | Infusion equipment for medical use Part 4: Infusion sets for single use, gravity feed
6. The characteristics and performance of a device shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the device, as indicated by the manufacturer, when the device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained in accordance with the manufacturer's instructions. | Y | ISO 14971 ISO 80369-7 ISO 80369-20 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 9626 | Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods
7. Devices shall be designed, manufactured and packaged in such a way that their characteristics and performance during their intended use are not adversely affected during transport and storage, for example, through fluctuations of temperature and humidity, taking account of the instructions and information provided by the manufacturer. | Y | ISO 14971 ISO 80369-7 ISO 80369-20 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | IEC 60601-1 | Medical electrical equipment - ALL PARTS
8. All known and foreseeable risks, and any undesirable side-effects, shall be minimised and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | IEC 60601-1-2 | Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests
9. For the devices referred to in Annex XVI, the general safety requirements set out in Sections 1 and 8 shall be understood to mean that the device, when used under the conditions and for the purposes intended, does not present a risk at all or presents a risk that is no more than the maximum acceptable risk related to the product's use which is consistent with a high level of protection for the safety and health of persons. | N | N/A | N/A | IEC 60601-1-6 | Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability
|  |  |  | IEC 62304 | Medical device software — Software life cycle processes
Chapter II — Requirements regarding design and manufacture | Apply Y/N | Relevant Standard(s) | Relevant device(s) | IEC 62366-1 | Medical devices Part 1: Application of usability engineering to medical devices
10. Chemical, physical and biological properties |  |  |  | ISO 10993-7 | Biological evaluation of medical devices Part 7: Ethylene oxide sterilization residuals
10.1. Devices shall be designed and manufactured in such a way as to ensure that the characteristics and performance requirements referred to in Chapter I are fulfilled. Particular attention shall be paid to: |  |  |  | ISO 11135 | Sterilization of health-care products — Ethylene oxide — Requirements for the development, validation and routine control of a sterilization process for medical devices
(a) the choice of materials and substances used, particularly as regards toxicity and, where relevant, flammability; | Y | ISO 14971 ISO 10555-1 ISO 10555-5 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-6 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 80369-7 ISO 80369-20 ISO 8536-4 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 14937 | Sterilization of health care products — General requirements for characterization of a sterilizing agent and the development, validation and routine control of a sterilization process for medical devices
(b) the compatibility between the materials and substances used and biological tissues, cells and body fluids, taking account of the intended purpose of the device and, where relevant, absorption, distribution, metabolism and excretion; | Y | ISO 10555-1 ISO 10555-5 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-6 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems | ISO 20417 | Medical devices — Information to be supplied by the manufacturer
(c) the compatibility between the different parts of a device which consists of more than one implantable part; | Y | ISO 10555-1 ISO 10555-5 ISO 10993-1 ISO 10993-3 ISO 10993-4 ISO 10993-5 ISO 10993-6 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) the impact of processes on material properties; | Y | ISO 13485 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(e) where appropriate, the results of biophysical or modelling research the validity of which has been demonstrated beforehand; | N | N/A | N/A
(f) the mechanical properties of the materials used, reflecting, where appropriate, matters such as strength, ductility, fracture resistance, wear resistance and fatigue resistance; | Y | ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(g) surface properties; and | Y | ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(h) the confirmation that the device meets any defined chemical and/or physical specifications. | Y | ISO 8536-4 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
10.2. Devices shall be designed, manufactured and packaged in such a way as to minimise the risk posed by contaminants and residues to patients, taking account of the intended purpose of the device, and to the persons involved in the transport, storage and use of the devices. Particular attention shall be paid to tissues exposed to those contaminants and residues and to the duration and frequency of exposure. | Y | ISO 14971 IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
10.3. Devices shall be designed and manufactured in such a way that they can be used safely with the materials and substances, including gases, with which they enter into contact during their intended use; if the devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned in accordance with the provisions and restrictions governing those medicinal products and that the performance of both the medicinal products and of the devices is maintained in accordance with their respective indications and intended use. | Y | ISO 13485 ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
10.4. Substances
10.4.1. Design and manufacture of devices
Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by substances or particles, including wear debris, degradation products and processing residues, that may be released from the device. Devices, or those parts thereof or those materials used therein that:
— are invasive and come into direct contact with the human body, | Y | ISO 14971 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— (re)administer medicines, body liquids or other substances, including gases, to/from the body, or | Y | ISO 14971 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— transport or store such medicines, body fluids or substances, including gases, to be (re)administered to the body, | Y | ISO 14971 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
shall only contain the following substances in a concentration that is above 0,1 % weight by weight (w/w) where justified pursuant to Section 10.4.2:
(a) substances which are carcinogenic, mutagenic or toxic to reproduction (‘CMR’), of category 1A or 1B, in accordance with Part 3 of Annex VI to Regulation (EC) No 1272/2008 of the European Parliament and of the Council, or | Y | ISO 14971 ISO 10993-7 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(b) substances having endocrine-disrupting properties for which there is scientific evidence of probable serious effects to human health and which are identified either in accordance with the procedure set out in Article 59 of Regulation (EC) No 1907/2006 of the European Parliament and of the Council or, once a delegated act has been adopted by the Commission pursuant to the first subparagraph of Article 5(3) of Regulation (EU) No 528/2012 of the European Parliament and the Council, in accordance with the criteria that are relevant to human health amongst the criteria established therein. | N | N/A | N/A
10.4.2. Justification regarding the presence of CMR and/or endocrine-disrupting substances
The justification for the presence of such substances shall be based upon:
(a) an analysis and estimation of potential patient or user exposure to the substance; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(b) an analysis of possible alternative substances, materials or designs, including, where available, information about independent research, peer-reviewed studies, scientific opinions from relevant scientific committees and an analysis of the availability of such alternatives; |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(c) argumentation as to why possible substance and/ or material substitutes, if available, or design changes, if feasible, are inappropriate in relation to maintaining the functionality, performance and the benefit-risk ratios of the product; including taking into account if the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials; and |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
(d) where applicable and available, the latest relevant scientific committee guidelines in accordance with Sections 10.4.3. and 10.4.4. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
10.4.3. Guidelines on phthalates
For the purposes of Section 10.4., the Commission shall, as soon as possible and by 26 May 2018, provide the relevant scientific committee with a mandate to prepare guidelines that shall be ready before 26 May 2020. The mandate for the committee shall encompass at least a benefit-risk assessment of the presence of phthalates which belong to either of the groups of substances referred to in points (a) and (b) of Section 10.4.1. The benefit-risk assessment shall take into account the intended purpose and context of the use of the device, as well as any available alternative substances and alternative materials, designs or medical treatments. When deemed appropriate on the basis of the latest scientific evidence, but at least every five years, the guidelines shall be updated. | N | N/A | N/A
10.4.4. Guidelines on other CMR and endocrine-disrupting substances
Subsequently, the Commission shall mandate the relevant scientific committee to prepare guidelines as referred to in Section 10.4.3. also for other substances referred to in points (a) and (b) of Section 10.4.1., where appropriate. |  | Whether compliance with the requirement depends on the device's detailed situation; please confirm the intended purpose statement of the medical device(s).
10.4.5. Labelling
Where devices, parts thereof or materials used therein as referred to in Section 10.4.1. contain substances referred to in points (a) or (b) of Section 10.4.1. in a concentration above 0,1 % weight by weight (w/w), the presence of those substances shall be labelled on the device itself and/or on the packaging for each unit or, where appropriate, on the sales packaging, with the list of such substances. If the intended use of such devices includes treatment of children or treatment of pregnant or breastfeeding women or treatment of other patient groups considered particularly vulnerable to such substances and/or materials, information on residual risks for those patient groups and, if applicable, on appropriate precautionary measures shall be given in the instructions for use. | Y | ISO 14971 ISO 10993-7 | C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
10.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks posed by the unintentional ingress of substances into the device taking into account the device and the nature of the environment in which it is intended to be used. | Y | ISO 14971 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
10.6. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks linked to the size and the properties of particles which are or can be released into the patient's or user's body, unless they come into contact with intact skin only. Special attention shall be given to nanomaterials. | Y | ISO 14971 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
11. Infection and microbial contamination
11.1. Devices and their manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as possible the risk of infection to patients, users and, where applicable, other persons. The design shall:
(a) reduce as far as possible and appropriate the risks from unintended cuts and pricks, such as needle stick injuries, | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(b) allow easy and safe handling, | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(c) reduce as far as possible any microbial leakage from the device and/or microbial exposure during use, and | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) prevent microbial contamination of the device or its content such as specimens or fluids. | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
11.2. Where necessary devices shall be designed to facilitate their safe cleaning, disinfection, and/or re-sterilisation. | Y | IEC 60601-1 IEC 60601-1-2 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
11.3. Devices labelled as having a specific microbial state shall be designed, manufactured and packaged to ensure that they remain in that state when placed on the market and remain so under the transport and storage conditions specified by the manufacturer. | N | N/A | N/A
11.4. Devices delivered in a sterile state shall be designed, manufactured and packaged in accordance with appropriate procedures, to ensure that they are sterile when placed on the market and that, unless the packaging which is intended to maintain their sterile condition is damaged, they remain sterile, under the transport and storage conditions specified by the manufacturer, until that packaging is opened at the point of use. It shall be ensured that the integrity of that packaging is clearly evident to the final user. | Y | ISO 13485 ISO 14971 IEC 60601-1-6 IEC 62304 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
11.5. Devices labelled as sterile shall be processed, manufactured, packaged and, sterilised by means of appropriate, validated methods. | Y | IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 62304 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
11.6. Devices intended to be sterilised shall be manufactured and packaged in appropriate and controlled conditions and facilities. | Y | IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 62304 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
11.7. Packaging systems for non-sterile devices shall maintain the integrity and cleanliness of the product and, where the devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the manufacturer. | Y | ISO 13485 ISO 14971 IEC 62366-1
11.8. The labelling of the device shall distinguish between identical or similar devices placed on the market in both a sterile and a non-sterile condition additional to the symbol used to indicate that devices are sterile. | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
12. Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body.
12.1. In the case of devices referred to in the first subparagraph of Article 1(8), the quality, safety and usefulness of the substance which, if used separately, would be considered to be a medicinal product within the meaning of point (2) of Article 1 of Directive 2001/83/EC, shall be verified by analogy with the methods specified in Annex I to Directive 2001/83/EC, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
12.2. Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body, and that are absorbed by or locally dispersed in the human body shall comply, where applicable and in a manner limited to the aspects not covered by this Regulation, with the relevant requirements laid down in Annex I to Directive 2001/83/EC for the evaluation of absorption, distribution, metabolism, excretion, local tolerance, toxicity, interaction with other devices, medicinal products or other substances and potential for adverse reactions, as required by the applicable conformity assessment procedure under this Regulation. | N | N/A | N/A
13. Devices incorporating materials of biological origin
13.1. For devices manufactured utilising derivatives of tissues or cells of human origin which are non-viable or are rendered non-viable covered by this Regulation in accordance with point (g) of Article 1(6), the following shall apply:
(a) donation, procurement and testing of the tissues and cells shall be done in accordance with Directive 2004/23/EC; | N | N/A | N/A
(b) processing, preservation and any other handling of those tissues and cells or their derivatives shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process; | N | N/A | N/A
(c) the traceability system for those devices shall be complementary and compatible with the traceability and data protection requirements laid down in Directive 2004/23/EC and in Directive 2002/98/EC. | N | N/A | N/A
13.2. For devices manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable the following shall apply:
(a) where feasible taking into account the animal species, tissues and cells of animal origin, or their derivatives, shall originate from animals that have been subjected to veterinary controls that are adapted to the intended use of the tissues. Information on the geographical origin of the animals shall be retained by manufacturers; | N | N/A | N/A
(b) sourcing, processing, preservation, testing and handling of tissues, cells and substances of animal origin, or their derivatives, shall be carried out so as to provide safety for patients, users and, where applicable, other persons. In particular safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or viral inactivation in the course of the manufacturing process, except when the use of such methods would lead to unacceptable degradation compromising the clinical benefit of the device; | N | N/A | N/A
(c) in the case of devices manufactured utilising tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012 the particular requirements laid down in that Regulation shall apply. | N | N/A | N/A
13.3. For devices manufactured utilising non-viable biological substances other than those referred to in Sections 13.1 and 13.2, the processing, preservation, testing and handling of those substances shall be carried out so as to provide safety for patients, users and, where applicable, other persons, including in the waste disposal chain. In particular, safety with regard to viruses and other transmissible agents shall be addressed by appropriate methods of sourcing and by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. | N | N/A | N/A
14. Construction of devices and interaction with their environment
14.1. If the device is intended for use in combination with other devices or equipment the whole combination, including the connection system shall be safe and shall not impair the specified performance of the devices. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. Connections which the user has to handle, such as fluid, gas transfer, electrical or mechanical coupling, shall be designed and constructed in such a way as to minimise all possible risks, such as misconnection. | Y | ISO 14971 ISO 23908 ISO 7864 ISO 80369-7 ISO 80369-20 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
14.2. Devices shall be designed and manufactured in such a way as to remove or reduce as far as possible:
(a) the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; | Y | ISO 14971 ISO 10993-10 ISO 10993-11 ISO 10993-12 ISO 14937 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(b) risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, temperature, variations in pressure and acceleration or radio signal interferences; | Y | ISO 14971 ISO 10555-1 ISO 80369-7 ISO 80369-20 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(c) the risks associated with the use of the device when it comes into contact with materials, liquids, and substances, including gases, to which it is exposed during normal conditions of use; | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) the risks associated with the possible negative interaction between software and the IT environment within which it operates and interacts; | Y | ISO 14971 ISO 80369-7 ISO 80369-20 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(e) the risks of accidental ingress of substances into the device; | Y | ISO 14971 ISO 10555-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(f) the risks of reciprocal interference with other devices normally used in the investigations or for the treatment given; and | Y | ISO 14971 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(g) risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Y | ISO 14971 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
14.3. Devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to devices the intended use of which includes exposure to or use in association with flammable or explosive substances or substances which could cause combustion. | Y | ISO 13485 ISO 14971 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
14.4. Devices shall be designed and manufactured in such a way that adjustment, calibration, and maintenance can be done safely and effectively. | Y | ISO 13485 ISO 14971 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
14.5. Devices that are intended to be operated together with other devices or products shall be designed and manufactured in such a way that the interoperability and compatibility are reliable and safe. | Y | ISO 14971 ISO 10555-1 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
14.6. Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking account of the intended purpose, users and the environmental conditions in which the devices are intended to be used. | N | N/A | N/A
14.7. Devices shall be designed and manufactured in such a way as to facilitate their safe disposal and the safe disposal of related waste substances by the user, patient or other person. To that end, manufacturers shall identify and test procedures and measures as a result of which their devices can be safely disposed after use. Such procedures shall be described in the instructions for use. | Y | ISO 13485 ISO 14971 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
15. Devices with a diagnostic or measuring function
15.1. Diagnostic devices and devices with a measuring function, shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. The limits of accuracy shall be indicated by the manufacturer. | N | N/A | N/A
15.2. The measurements made by devices with a measuring function shall be expressed in legal units conforming to the provisions of Council Directive 80/181/EEC. | N | N/A | N/A
16. Protection against radiation
16.1. General
(a) Devices shall be designed, manufactured and packaged in such a way that exposure of patients, users and other persons to radiation is reduced as far as possible, and in a manner that is compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes. | N | N/A | N/A
(b) The operating instructions for devices emitting hazardous or potentially hazardous radiation shall contain detailed information as to the nature of the emitted radiation, the means of protecting the patient and the user, and on ways of avoiding misuse and of reducing the risks inherent to installation as far as possible and appropriate. Information regarding the acceptance and performance testing, the acceptance criteria, and the maintenance procedure shall also be specified. | N | N/A | N/A
16.2. Intended radiation
(a) Where devices are designed to emit hazardous, or potentially hazardous, levels of ionizing and/or non-ionizing radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent to the emission, it shall be possible for the user to control the emissions. Such devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | N | N/A | N/A
(b) Where devices are intended to emit hazardous, or potentially hazardous, ionizing and/or non-ionizing radiation, they shall be fitted, where possible, with visual displays and/or audible warnings of such emissions. | N | N/A | N/A
16.3. Devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as possible. Where possible and appropriate, methods shall be selected which reduce the exposure to radiation of patients, users and other persons who may be affected. | N | N/A | N/A
16.4. Ionising radiation
(a) Devices intended to emit ionizing radiation shall be designed and manufactured taking into account the requirements of the Directive 2013/59/Euratom laying down basic safety standards for protection against the dangers arising from exposure to ionising radiation. | N | N/A | N/A
(b) Devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where possible, taking into account the intended use, the quantity, geometry and quality of the radiation emitted can be varied and controlled, and, if possible, monitored during treatment. | N | N/A | N/A
(c) Devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve an image and/or output quality that are appropriate to the intended medical purpose whilst minimising radiation exposure of the patient and user. | N | N/A | N/A
(d) Devices that emit ionising radiation and are intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type, energy and, where appropriate, the quality of radiation. | N | N/A | N/A
17. Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves
17.1. Devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, shall be designed to ensure repeatability, reliability and performance in line with their intended use. In the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks or impairment of performance. | N | N/A | N/A
17.2. For devices that incorporate software or for software that are devices in themselves, the software shall be developed and manufactured in accordance with the state of the art taking into account the principles of development life cycle, risk management, including information security, verification and validation. | N | N/A | N/A
17.3. Software referred to in this Section that is intended to be used in combination with mobile computing platforms shall be designed and manufactured taking into account the specific features of the mobile platform (e.g. size and contrast ratio of the screen) and the external factors related to their use (varying environment as regards level of light or noise). | N | N/A | N/A
17.4. Manufacturers shall set out minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | N | N/A | N/A
18. Active devices and devices connected to them
18.1. For non-implantable active devices, in the event of a single fault condition, appropriate means shall be adopted to eliminate or reduce as far as possible consequent risks. | N | N/A | N/A
18.2. Devices where the safety of the patient depends on an internal power supply shall be equipped with a means of determining the state of the power supply and an appropriate warning or indication for when the capacity of the power supply becomes critical. If necessary, such warning or indication shall be given prior to the power supply becoming critical. | N | N/A | N/A
18.3. Devices where the safety of the patient depends on an external power supply shall include an alarm system to signal any power failure. | N | N/A | N/A
18.4. Devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health. | N | N/A | N/A
18.5. Devices shall be designed and manufactured in such a way as to reduce as far as possible the risks of creating electromagnetic interference which could impair the operation of the device in question or other devices or equipment in the intended environment. | N | N/A | N/A
18.6. Devices shall be designed and manufactured in such a way as to provide a level of intrinsic immunity to electromagnetic interference such that is adequate to enable them to operate as intended. | N | N/A | N/A
18.7. Devices shall be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks to the patient, user or any other person, both during normal use of the device and in the event of a single fault condition in the device, provided the device is installed and maintained as indicated by the manufacturer. | N | N/A | N/A
18.8. Devices shall be designed and manufactured in such a way as to protect, as far as possible, against unauthorised access that could hamper the device from functioning as intended. | N | N/A | N/A
19. Particular requirements for active implantable devices
19.1. Active implantable devices shall be designed and manufactured in such a way as to remove or minimize as far as possible:
(a) risks connected with the use of energy sources with particular reference, where electricity is used, to insulation, leakage currents and overheating of the devices, | N | N/A | N/A
(b) risks connected with medical treatment, in particular those resulting from the use of defibrillators or high-frequency surgical equipment, and | N | N/A | N/A
(c) risks which may arise where maintenance and calibration are impossible, including:
— excessive increase of leakage currents, | N | N/A | N/A
— ageing of the materials used, | N | N/A | N/A
— excess heat generated by the device, | N | N/A | N/A
— decreased accuracy of any measuring or control mechanism. | N | N/A | N/A
19.2. Active implantable devices shall be designed and manufactured in such a way as to ensure
— if applicable, the compatibility of the devices with the substances they are intended to administer, and | N | N/A | N/A
— the reliability of the source of energy. | N | N/A | N/A
19.3. Active implantable devices and, if appropriate, their component parts shall be identifiable to allow any necessary measure to be taken following the discovery of a potential risk in connection with the devices or their component parts. | N | N/A | N/A
19.4. Active implantable devices shall bear a code by which they and their manufacturer can be unequivocally identified (particularly with regard to the type of device and its year of manufacture); it shall be possible to read this code, if necessary, without the need for a surgical operation. | N | N/A | N/A
20. Protection against mechanical and thermal risks
20.1. Devices shall be designed and manufactured in such a way as to protect patients and users against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | Y | ISO 14971 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
20.2. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. | Y | ISO 14971 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
20.3. Devices shall be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | N | N/A | N/A
20.4. Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user or other person has to handle, shall be designed and constructed in such a way as to minimise all possible risks. | Y | ISO 14971 ISO 80369-7 ISO 80369-20 ISO 8536-4 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
20.5. Errors likely to be made when fitting or refitting certain parts which could be a source of risk shall be made impossible by the design and construction of such parts or, failing this, by information given on the parts themselves and/or their housings. | Y | ISO 14971 IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
The same information shall be given on moving parts and/or their housings where the direction of movement needs to be known in order to avoid a risk. | Y | ISO 14971 IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
20.6. Accessible parts of devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal conditions of use. | Y | ISO 14971 ISO 8536-4 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
21. Protection against the risks posed to the patient or user by devices supplying energy or substances
21.1. Devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the amount to be delivered can be set and maintained accurately enough to ensure the safety of the patient and of the user. | Y | ISO 10993-10 ISO 10993-11 ISO 23908 ISO 7864 ISO 9626 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010502 - Intravascular emobilic protection catheters and systems
21.2. Devices shall be fitted with the means of preventing and/or indicating any inadequacies in the amount of energy delivered or substances delivered which could pose a danger. Devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy or substances from an energy and/or substance source. | Y | ISO 14971 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010502 - Intravascular emobilic protection catheters and systems
21.3. The function of the controls and indicators shall be clearly specified on the devices. Where a device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient. | Y | ISO 8536-4 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
22. Protection against the risks posed by medical devices intended by the manufacturer for use by lay persons
22.1. Devices for use by lay persons shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can be reasonably anticipated in the lay person's technique and environment. The information and instructions provided by the manufacturer shall be easy for the lay person to understand and apply. | N | N/A | N/A
22.2. Devices for use by lay persons shall be designed and manufactured in such a way as to:
— ensure that the device can be used safely and accurately by the intended user at all stages of the procedure, if necessary after appropriate training and/or information, | N | N/A | N/A
— reduce, as far as possible and appropriate, the risk from unintended cuts and pricks such as needle stick injuries, and | N | N/A | N/A
— reduce as far as possible the risk of error by the intended user in the handling of the device and, if applicable, in the interpretation of the results. | N | N/A | N/A
22.3. Devices for use by lay persons shall, where appropriate, include a procedure by which the lay person:
— can verify that, at the time of use, the device will perform as intended by the manufacturer, and | N | N/A | N/A
— if applicable, is warned if the device has failed to provide a valid result. | N | N/A | N/A
Chapter III — Requirements regarding the information supplied with the device | Apply Y/N | Relevant Standard(s) | Relevant device(s)
23. Label and instructions for use
23.1. General requirements regarding the information supplied by the manufacturer
Each device shall be accompanied by the information needed to identify the device and its manufacturer, and by any safety and performance information relevant to the user, or any other person, as appropriate. Such information may appear on the device itself, on the packaging or in the instructions for use, and shall, if the manufacturer has a website, be made available and kept up to date on the website, taking into account the following:
(a) The medium, format, content, legibility, and location of the label and instructions for use shall be appropriate to the particular device, its intended purpose and the technical knowledge, experience, education or training of the intended user(s). In particular, instructions for use shall be written in terms readily understood by the intended user and, where appropriate, supplemented with drawings and diagrams. | Y | ISO 8536-4 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(b) The information required on the label shall be provided on the device itself. If this is not practicable or appropriate, some or all of the information may appear on the packaging for each unit, and/or on the packaging of multiple devices. | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(c) Labels shall be provided in a human-readable format and may be supplemented by machine-readable information, such as radio-frequency identification (‘RFID’) or bar codes. | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) Instructions for use shall be provided together with devices. By way of exception, instructions for use shall not be required for class I and class IIa devices if such devices can be used safely without any such instructions and unless otherwise provided for elsewhere in this Section. | N | N/A | N/A
(e) Where multiple devices are supplied to a single user and/or location, a single copy of the instructions for use may be provided if so agreed by the purchaser who in any case may request further copies to be provided free of charge. | N | N/A | N/A
(f) Instructions for use may be provided to the user in non-paper format (e.g. electronic) to the extent, and only under the conditions, set out in Regulation (EU) No 207/2012 or in any subsequent implementing rules adopted pursuant to this Regulation. | N | N/A | N/A
(g) Residual risks which are required to be communicated to the user and/or other person shall be included as limitations, contra-indications, precautions or warnings in the information supplied by the manufacturer. | Y | ISO 14971 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(h) Where appropriate, the information supplied by the manufacturer shall take the form of internationally recognised symbols. Any symbol or identification colour used shall conform to the harmonised standards or CS. In areas for which no harmonised standards or CS exist, the symbols and colours shall be described in the documentation supplied with the device. | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
23.2. Information on the label
The label shall bear all of the following particulars:
(a) the name or trade name of the device; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(b) the details strictly necessary for a user to identify the device, the contents of the packaging and, where it is not obvious for the user, the intended purpose of the device; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(c) the name, registered trade name or registered trade mark of the manufacturer and the address of its registered place of business; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) if the manufacturer has its registered place of business outside the Union, the name of the authorised representative and address of the registered place of business of the authorised representative; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(e) where applicable, an indication that the device contains or incorporates:
— a medicinal substance, including a human blood or plasma derivative, or | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— tissues or cells, or their derivatives, of human origin, or | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— tissues or cells of animal origin, or their derivatives, as referred to in Regulation (EU) No 722/2012; | N | N/A | N/A
(f) where applicable, information labelled in accordance with Section 10.4.5.; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(g) the lot number or the serial number of the device preceded by the words LOT NUMBER or SERIAL NUMBER or an equivalent symbol, as appropriate; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(h) the UDI carrier referred to in Article 27(4) and Part C of Annex VII; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(i) an unambiguous indication of t the time limit for using or implanting the device safely, expressed at least in terms of year and month, where this is relevant; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(j) where there is no indication of the date until when it may be used safely, the date of manufacture. This date of manufacture may be included as part of the lot number or serial number, provided the date is clearly identifiable; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(k) an indication of any special storage and/or handling condition that applies; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(l) if the device is supplied sterile, an indication of its sterile state and the sterilisation method; | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(m) warnings or precautions to be taken that need to be brought to the immediate attention of the user of the device, and to any other person. This information may be kept to a minimum in which case more detailed information shall appear in the instructions for use, taking into account the intended users; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(n) if the device is intended for single use, an indication of that fact. A manufacturer's indication of single use shall be consistent across the Union; | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(o) if the device is a single-use device that has been reprocessed, an indication of that fact, the number of reprocessing cycles already performed, and any limitation as regards the number of reprocessing cycles; | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(p) if the device is custom-made, the words ‘custom-made device’; | Y | ISO 8536-4 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(q) an indication that the device is a medical device. If the device is intended for clinical investigation only, the words ‘exclusively for clinical investigation’; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(r) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body, the overall qualitative composition of the device and quantitative information on the main constituent or constituents responsible for achieving the principal intended action; | Y | ISO 8536-4 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(s) for active implantable devices, the serial number, and for other implantable devices, the serial number or the lot number. | N | N/A | N/A
23.3. Information on the packaging which maintains the sterile condition of a device (‘sterile packaging’)
The following particulars shall appear on the sterile packaging:
(a) an indication permitting the sterile packaging to be recognised as such, | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(b) a declaration that the device is in a sterile condition, | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(c) the method of sterilisation, | Y | IEC 60601-1 IEC 60601-1-2 IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) the name and address of the manufacturer, | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(e) a description of the device, | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(f) if the device is intended for clinical investigations, the words ‘exclusively for clinical investigations’, | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(g) if the device is custom-made, the words ‘custom-made device’, | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(h) the month and year of manufacture, | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(i) an unambiguous indication of the time limit for using or implanting the device safely expressed at least in terms of year and month, and | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(j) an instruction to check the instructions for use for what to do if the sterile packaging is damaged or unintentionally opened before use. | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
23.4. Information in the instructions for use
The instructions for use shall contain all of the following particulars:
(a) the particulars referred to in points (a), (c), (e), (f), (k), (l), (n) and (r) of Section 23.2; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(b) the device's intended purpose with a clear specification of indications, contra-indications, the patient target group or groups, and of the intended users, as appropriate; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(c) where applicable, a specification of the clinical benefits to be expected. | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(d) where applicable, links to the summary of safety and clinical performance referred to in Article 32; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(e) the performance characteristics of the device; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(f) where applicable, information allowing the healthcare professional to verify if the device is suitable and select the corresponding software and accessories; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(g) any residual risks, contra-indications and any undesirable side-effects, including information to be conveyed to the patient in this regard; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(h) specifications the user requires to use the device appropriately, e.g. if the device has a measuring function, the degree of accuracy claimed for it; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(i) details of any preparatory treatment or handling of the device before it is ready for use or during its use, such as sterilisation, final assembly, calibration, etc., including the levels of disinfection required to ensure patient safety and all available methods for achieving those levels of disinfection; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(j) any requirements for special facilities, or special training, or particular qualifications of the device user and/or other persons; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(k) the information needed to verify whether the device is properly installed and is ready to perform safely and as intended by the manufacturer, together with, where relevant:
— details of the nature, and frequency, of preventive and regular maintenance, and of any preparatory cleaning or disinfection, | Y | IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— identification of any consumable components and how to replace them, | Y | IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— information on any necessary calibration to ensure that the device operates properly and safely during its intended lifetime, and | Y | IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— methods for eliminating the risks encountered by persons involved in installing, calibrating or servicing devices; | Y | IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(l) if the device is supplied sterile, instructions in the event of the sterile packaging being damaged or unintentionally opened before use; | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(m) if the device is supplied non-sterile with the intention that it is sterilised before use, the appropriate instructions for sterilisation; | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(n) if the device is reusable, information on the appropriate processes for allowing reuse, including cleaning, disinfection, packaging and, where appropriate, the validated method of re-sterilisation appropriate to the Member State or Member States in which the device has been placed on the market. Information shall be provided to identify when the device should no longer be reused, e.g. signs of material degradation or the maximum number of allowable reuses; | Y | IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(o) an indication, if appropriate, that a device can be reused only if it is reconditioned under the responsibility of the manufacturer to comply with the general safety and performance requirements; | N | N/A | N/A
(p) if the device bears an indication that it is for single use, information on known characteristics and technical factors known to the manufacturer that could pose a risk if the device were to be re-used. This information shall be based on a specific section of the manufacturer's risk management documentation, where such characteristics and technical factors shall be addressed in detail. If in accordance with point (d) of Section 23.1. no instructions for use are required, this information shall be made available to the user upon request; | Y | ISO 14971 IEC 60601-1-6 IEC 62304 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(q) for devices intended for use together with other devices and/or general purpose equipment:
— information to identify such devices or equipment, in order to obtain a safe combination, and/or | Y | ISO 8536-4 IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— information on any known restrictions to combinations of devices and equipment; | Y | ISO 8536-4 IEC 62366-1 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(r) if the device emits radiation for medical purposes:
— detailed information as to the nature, type and where appropriate, the intensity and distribution of the emitted radiation, | Y | IEC 62366-1 ISO 10993-7 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— the means of protecting the patient, user, or other person from unintended radiation during use of the device; | Y | IEC 62366-1 ISO 10993-7 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(s) information that allows the user and/or patient to be informed of any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. That information shall, where relevant, allow the user to brief the patient about any warnings, precautions, contra-indications, measures to be taken and limitations of use regarding the device. The information shall cover, where appropriate:
— warnings, precautions and/or measures to be taken in the event of malfunction of the device or changes in its performance that may affect safety, | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— warnings, precautions and/or measures to be taken as regards the exposure to reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, radiation associated with diagnostic or therapeutic procedures, pressure, humidity, or temperature, | Y | IEC 62366-1 ISO 10993-7 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— warnings, precautions and/or measures to be taken as regards the risks of interference posed by the reasonably foreseeable presence of the device during specific diagnostic investigations, evaluations, or therapeutic treatment or other procedures such as electromagnetic interference emitted by the device affecting other equipment, | Y | IEC 62366-1 ISO 10993-7 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
— if the device is intended to administer medicinal products, tissues or cells of human or animal origin, or their derivatives, or biological substances, any limitations or incompatibility in the choice of substances to be delivered, | Y | ISO 10555-1 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— warnings, precautions and/or limitations related to the medicinal substance or biological material that is incorporated into the device as an integral part of the device; and | Y | ISO 10555-1 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— precautions related to materials incorporated into the device that contain or consist of CMR substances or endocrine-disrupting substances, or that could result in sensitisation or an allergic reaction by the patient or user; | Y | ISO 14971 ISO 10993-7 | C010103 - Integrated peripheral I.V. devices C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters
(t) in the case of devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body and that are absorbed by or locally dispersed in the human body, warnings and precautions, where appropriate, related to the general profile of interaction of the device and its products of metabolism with other devices, medicinal products and other substances as well as contra-indications, undesirable side-effects and risks relating to overdose; | Y | ISO 10555-1 IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(u) in the case of implantable devices, the overall qualitative and quantitative information on the materials and substances to which patients can be exposed; | N | N/A | N/A
(v) warnings or precautions to be taken in order to facilitate the safe disposal of the device, its accessories and the consumables used with it, if any. This information shall cover, where appropriate:
— infection or microbial hazards such as explants, needles or surgical equipment contaminated with potentially infectious substances of human origin, and | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
— physical hazards such as from sharps. | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
If in accordance with the point (d) of Section 23.1 no instructions for use are required, this information shall be made available to the user upon request; | N | N/A | N/A
(w) for devices intended for use by lay persons, the circumstances in which the user should consult a healthcare professional; | N | N/A | N/A
(x) for the devices covered by this Regulation pursuant to Article 1(2), information regarding the absence of a clinical benefit and the risks related to use of the device; | N | N/A | N/A
(y) date of issue of the instructions for use or, if they have been revised, date of issue and identifier of the latest revision of the instructions for use; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(z) a notice to the user and/or patient that any serious incident that has occurred in relation to the device should be reported to the manufacturer and the competent authority of the Member State in which the user and/or patient is established; | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems
(aa) information to be supplied to the patient with an implanted device in accordance with Article 18; | N | N/A | N/A
(ab) for devices that incorporate electronic programmable systems, including software, or software that are devices in themselves, minimum requirements concerning hardware, IT networks characteristics and IT security measures, including protection against unauthorised access, necessary to run the software as intended. | Y | IEC 62366-1 | C010101 - Peripheral I.V. catheters C010101 - Peripheral I.V. catheters C010103 - Integrated peripheral I.V. devices C010104 - Umbilical I.V. cannulas C010201 - Central I.V. catheters, peripheral access C010202 - Central I.V. catheters, non-tunneled C010203 - Central venous catheters, partially tunneled C010204 - Subcutaneous implantable venous access port systems C010301 - Radial cannulas C010302 - Umbilical cannulas C010303 - Femoral cannulas C010401 - Cardiac angiography devices C010402 - Peripheral angiography devices C010403 - Fractional flow reserve (ffr) measurement devices C010501 - Vena cava filters C010502 - Intravascular emobilic protection catheters and systems